## Slide 1: Radiation Dose Management in Computed Tomography
L04
Tube Potential in CT
Material of the e-learning program
IAEA, 2017

## Slide 2: Objectives
Understand basics of KV in CT
Understand effect of KV on dose and quality
Importance of KV in modern MDCT
Selection of appropriate kV
Automatic kV selection technique

## Slide 3: Tube potential (kV)
Represents potential difference between cathode and anode, which drives electrons across x-ray tube
Unit is kilovoltage (kV)
KV change has a profound effect on radiation dose and noise
Radiation dose ∝ Tube current
Radiation dose ∝ (kV)2

## Slide 4: kV and Radiation
Compared to 80 kV,
  X-ray output at 100 kV is 1.5 times higher
  X-ray output at 120 kV is 2.5 times higher
  X-ray output at 140 kV is 3.4 times higher
kV | Relative CTDI vol
80 kV | 0.4
100 kV | 0.7
120 kV | 1.0
140 kV | 1.4
Huda et al. Radiology 2000

## Slide 5: kV and Quality
kV | CT numbers (iodine) | Image noise | Radiation dose
kV decrease | Increase | Increase | Decrease
kV increase | Decrease | Decrease | Increase
80 kV
140 kV
Decrease in kV brings x-ray energy closer to the k-edge of iodine, which results in its increased x-ray attenuation and higher HU
For contrast enhanced CT: kV reduction improves image contrast while reducing dose

## Slide 6: Selecting the right KV
kV can be selected or specified manually
  Most common and widely available
kV can be automatically selected by CT
  Available from some vendors (Siemens & GE)
  Unlike AEC which varies mA, automatic kV selection technique does not modulate kV but only picks a best kV for the entire CT exam.
Most scanners allow 4 choices of kV
  80, 100, 120, 140 (some have 135 instead of 140)
Some scanners only have 3 available kV
  80, 120 and 140 kV
Some scanners have 5 available kV
  70, 80, 100, 120, 140 kV (some Siemens CT)
Some scanners have even more!!
  70, 80, 90, 100, 110, 120, 130, 140, 150 (Force)

## Slide 7: Tube potential (kV)
Most scanners require users to specify a fixed kV for CT protocols
  Generally, smaller patients (particularly children) should be scanned at low kV (70-100 kV)
  Generally, CT angiography or contrast enhanced CT can be scanned at low kV (80-100 kV)
Some CT have automatic kV selection where CT scanner picks up the right kV
  Care kV (Siemens)
  KV Assist (GE)

## Slide 8: kV effects
kV decrease# | kV increase#
Noise increases May need mA increase Newer CT have high mA limit Iterative reconstruction reduce noise Photon starvation and beam hardening artifacts possible if kV is too low if patient is too large Image iodine contrast improves Contrast volume may be reduced Radiation dose is reduced | Noise decreases Can use lower mA Older scanners have low mA limit FBP may need higher kV for some Photon starvation and beam hardening artifacts may decrease if kV is high if patient is small Image iodine contrast deteriorates Larger contrast volume may be needed Radiation dose is increased
# If other scan parameters are held constant

## Slide 9: KV and its effect on HU, Noise (SD) and Dose
140 kV
CTDI vol: 29 mGy
120 kV
24 mGy
100 kV
16 mGy
80 kV
8 mGy
Mean HU 120
SD: 13
Mean HU 138
SD: 14
Mean HU 162
SD: 31
Mean HU 210
SD: 43

## Slide 10: kV Adjustments
Lower KV (≤100 kV) | Higher kV (>100kV)
Most body CT in children (< about 80 kg) Most CTA of neck and chest (80-100 kg) Post-contrast chest CT (< 100 kg) Arterial phase CT of abdomen and abdominal CTA (up to 80 kg) Extremities CTA Very low dose CT: Lung nodule follow up Lung cancer screening CT colonography (80 kg) | Very large children ( > about 100 kg) Large shoulders - CTA of neck (>100 kg) Large patients (> 100 kg) for chest CTA angiography or post-contrast chest CT Portal venous phase CT of abdomen Abdominal CTA or arterial phase CT (>80 kg) Subtle lesions of the abdomen Low attenuation liver lesions Pancreatic cancer staging Solid renal tumor
Newer CT systems and iterative reconstruction techniques can enable use of lower kV.
More frequently and in larger patients compared to older scanners with FBP.
CTA = CT angiography

## Slide 11: kV and Head CT
Most often routine head CT in adults are performed at 120 kV
Most children can be scanned at ≤100 kV
CT angiography of head and neck should be performed at ≤ 100 kV (< 80 - 100 kg)
CT for ventricular shunt patency and craniosynostosis should be performed at lower kV
Avoid low kV in patients with large shoulders undergoing neck CT or CT angiography

## Slide 12: kV and Chest CT
Most chest CT can be scanned at ≤ 100 kV (in patients 80 – 100 kg)
Pulmonary embolism CT can be at 80 or 100 kV (in patients 80 – 100 kg)
Lung nodule follow up CT and lung cancer screening CT can also be performed at ≤100 kV
All pediatric chest CT can be performed at 80/100 kV (less than 80 kg)
When available 70 kV should be used for chest CT in infants
70 kV should be used for chest CT in infants

## Slide 13: CT pulmonary angiography at low kV
With reduction in kVp, radiation dose decreases while HU number & image contras increase
558 HU
278 HU
140 kV
80 kV

## Slide 14: kV and Cardiac CT
Contrast enhanced coronary arteries allow lower kV use
Most coronary CT angiography in patients less than 30 kg/m2 can be and should be performed at 80-100 kV
Most children should also be scanned at 80 kV for coronary CT angiography, since major concern is anomalous coronary arteries

## Slide 15: Cardiac CT and low kV
0.75 mm
2 mm
Age 14 years Weight 60 kg EKG gated CCTA Pitch 3.4 80 kV | Age 65 years Weight 140 kg Non gated Pulm. veins Pitch 3 100 kV

## Slide 16: kV and Abdominal CT
Most abdominal CT are performed at 120 kV
With iterative reconstruction techniques, many abdominal CT in non-obese patients are being performed at 100 kV
CT angiography and arterial phase abdominal CT with high iodine presence should be scanned at lower kV (≤100 kV) in most non-obese patients
Newer CT with higher mAs limit (800 - 1300 mAs) are expected to increase use of lower kV in abdomen

## Slide 17: Multiphase Liver CT: Dose Reduction
Arterial phase:
Lower kV (100 kV)
Liver only coverage
AEC
Portal venous phase:
Longer coverage
AEC

## Slide 18: kV, Dose and HU
kV: 140 120 (-35% dose) 100 (-50% dose)
HU: 300 357 470

## Slide 19: Automatic kV Selection
      System automatically identifies optimum kV setting (Care kV, Siemens)
        Body habitus (from planning radiograph)
        Exam type (non-contrast, bone, standard contrast, vascular)
        System limits (tube current)
      System automatically proposes kV and corresponding mAs values to maintain constant contrast to noise ratio.
      kV Assist (GE) also works on similar principles for selecting optimal kV
Notes: 19

## Slide 20: Summary
kV is a unique scan parameter that affects HU and radiation dose.
Use of lower kV scanning is increasing on newer generation CT systems and with iterative reconstruction enabled image noise reduction.
Lower kV should be used to reduce radiation dose in most children and CT angiography studies.