## Slide 1: Radiation Dose Management in Computed Tomography
L08
Key Aspects of Abdomen CT
Material of the e-learning program
IAEA, 2017

## Slide 2: Objectives
Understand key protocols based on clinical indications for abdominal CT
Understand scan parameters used for abdominal CT
Ensure proper use of multiple scanning passes in abdomen

## Slide 3: Is Abdomen unique for CT?
Abdomen CT is the most frequent CT in modern medicine
Certain diseases and organs allow substantial dose reduction due to high inherent contrast
  Kidney stones
  CT colonography
Certain abdominal organs have propensity for subtle or low contrast lesions
  Need lower noise images
  Need multiple passes through the same region

## Slide 4: Protocols: Meeting Basic Standards

## Slide 5: Developing CT protocols
“On the fly” protocol creation is not optimal
Valuable time can be wasted for urgent CT
Level of training between radiographers and between radiologists is different
Errors can result in implementation since there are far too many scan parameters and protocols
Consistency is difficult to achieve
  Same clinical indication and size: Similar doses
  Different clinical indications or size: Different doses

## Slide 6: Appropriate clinical reason
Does not matter chest or abdomen, head or big toe, the best way to tackle dose remains…

## Slide 7: Specific Abdominal CT Protocols
All scan protocols must begin with the clinical indications for their use
Scan protocols must be based on clinical indications
Each scan protocol should then address
  Number of scan phases required
  Scan range for each phase
  Scan parameters for each phase
  Dose adjustment for patient size

## Slide 8: Need: Indication driven protocols?
Certain things can be seen at reduced dose
  Urinary stones
  CT colonography
  CT enterography
Others need higher dose for assessment
  Low attenuation liver lesions
  Pancreatic neoplasms
  Solid renal tumors
CT Colonography < CT urinary calculi < CT routine abdomen < CT biphasic or triphasic liver protocol

## Slide 9: Abdomen CT protocols: Good Points
First, develop clinical indication specific protocols
Then, ensure dose adjustment to patient size
Tube Current:
  Perform abdominal CT with AEC
  Avoid use of fixed mA for abdominal CT
  Adjust mA for different clinical indications
    Lower mAs for urinary stones/CT colonography
    Lower mAs for iterative reconstructions than FBP
kV selection: Automated or Manual
  Use lower kV for small patients and children
  Use lower kV for CT angiography (arterial phase) unless large patient

## Slide 10: Abdomen CT protocols: Good Points
  Scanning mode:
    Helical acquisition for most abdomen CT
  Pitch for Abdominal CT
    Non-overlapping pitch should be preferred to minimize scan duration and motion artifacts
      Pitch should be ≥ 1:1
Beam collimation:
    Wider beam collimation for most routine CT: More dose efficient
Gantry rotation speed:
    Fast to minimize motion artifacts (0.4-0.5 seconds)
    May need to be slowed down for very large patients to increase mAs

## Slide 11: Abdomen CT protocols: Good Points
  Number of scan series for abdomen CT
  Unfortunately, repeated scanning is common in abdomen
  Routine pre-contrast prior to post-contrast CT should be avoided
  When performing multiple phases, question need and technique
Multiple phases | Questions to ask and answer
Need? | Routine non-contrast phase before post-contrast: No Routine arterial and venous phases: No Routine delayed images: No
Same length? | Length for some phases can be less: Yes Ex: arterial phase liver or pancreas can be smaller length Ex: Delayed phase: through the lesion only
Same dose? | Some phases can be acquired at lower dose: Yes Non-contrast: Lower dose Arterial phase: Lower KV to reduce dose Iterative reconstruction technique to enable dose reduction

## Slide 12: Scan Length: Abdominal CT
Indication | Coverage
Routine or R/O | Dome of liver - pubic symphysis
Delays | Through lesion only (not entire organ)
Kidney stone CT Urography | Top of kidneys - Symphysis
Dual phase liver | Arterial: Liver Portal venous: Entire abdomen
Appendix in young patients | Limited coverage: L3 to symphysis
Chest-Abdomen CT
Minimize scan overlap

## Slide 13: Iterative Reconstruction (IR) in Abdomen
Several studies have shown reduced radiation doses with IR in abdomen CT
Extent of dose reduction with IR relative to filtered back projection varies by clinical indication and patient size
Generally speaking, 30-50% dose reduction can be anticipated versus FBP for abdominal CT

## Slide 14: Select appropriate slice thickness for abdomen CT
Slice thickness 0.75 mm
All images were acquired at 13 mGy CTDI vol: Thinner slices have more noise but better contrast
Slice thickness 1.5 mm
Slice thickness 3 mm
Slice thickness 5 mm
Slice thickness 0.75 mm
Slice thickness 1.5 mm
Slice thickness 3 mm
Slice thickness 5 mm
All images were acquired at 3.5 mGy CTDI vol: Thinner slices have more noise but better contrast

## Slide 15: Iterative reconstruction (ASIR GE) help reduce image noise in low dose CT
ASIR70 - 12 mGy
ASIR70 - 9 mGy
ASIR70 - 6 mGy
ASIR70 - 3 mGy
ASIR50 - 12 mGy
ASIR50 - 9 mGy
ASIR50 - 6 mGy
ASIR50 - 3 mGy
ASIR30 - 12 mGy
ASIR30 - 9 mGy
ASIR30 - 6 mGy
ASIR30 - 3 mGy
FBP- 12 mGy
FBP- 9 mGy
FBP- 6 mGy
FBP -3 mGy

## Slide 16: Iterative reconstruction (IRIS, Siemens) help reduce image noise in low dose CT
IRIS 200 mAs
FBP 200 mAs
FBP 100 mAs
IRIS 100 mAs
FBP 50 mAs
IRIS 50 mAs

## Slide 17: Iterative reconstruction help reduce image noise in low dose CT
FBP 200 mA
FBP 160 mA
FBP 80 mA
S1 80 mA
S2 80 mA
S3 80 mA
S1, S2, and S3 represent increasing strength of Safire (Siemens) for noise reduction

## Slide 18: Indication Driven Abdomen CT Protocols
Protocol | Clinical Reasons | Specific instructions
Routine Abdomen | Masses, infections, pain, cancer staging (non-abdominal primary) | No routine non-contrast before contrast
CT urinary calculi | Suspected or known renal colic | Follow up stone CT at lower dose than initial
CT hematuria | < 40 years: non-contrast CT- Stone – No post contrast | > 40 years: non-contrast & post-contrast
CT adrenal protocol | Characterize adrenal nodule seen on chest or routine abdomen | All phases through adrenal region only
CT colonography | Screening exam, completion colonography | Lowest dose abdominal CT
CT biphasic liver | When MR can not be performed in patients with suspected liver malignancies | Arterial phase: Lower kV Arterial phase: Liver only Portal venous phase: entire abdomen

## Slide 19: CT Colonography
30-40 mAs: 100-120 kV
Flat polyp (8 mm)
2-4 mGy
High tissue contrast between air and colonic wall and lesions enable dose reduction

## Slide 20: What else is high contrast in Abdomen?
High contrast = Kidney stones Calcium vs Soft tissues
Lower dose

## Slide 21: Kidney stone CT: Seen at Lower Dose
FBP 300 mAs 140 kV 30 mGy
200 mAs 120 kV 13 mGy
100 mAs 120 kV 7 mGy
50 mAs 120 kV 3.4 mGy
350 mAs 120 kV 24 mGy
65 kg
BMI 26

## Slide 22: Urinary Stone CT
CT images of 60-year-old man acquired with fixed current (A) and
AEC show a tiny calculus (arrow) in left renal pelvis. AEC (B) enabled 50% dose reduction compared to fixed current technique.
Tube Potential | 120 KV commonly (100 kV< 60 kg)
Tube current (prefer AEC) | About 30-50% lower than routine abdomen
Image thickness (mm) | 2.5 - 5mm
CTDI vol | 2-6 mGy (size based)
A
B
Fixed mA
Use AEC (NI 20)

## Slide 23: Calcium : Seen at Lower Dose
Radiodense stones (GB) can also be seen at low dose
Note “fading” ascites with decreasing DOSE
FBP 300 mAs 140 kV 30 mGy
200 mAs 120 kV 13 mGy
350 mAs 120 kV 24 mGy
100 mAs 120 kV 7 mGy
50 mAs 120 kV 3.4 mGy

## Slide 24: Bony findings at low doses
200 mAs 12 mGy FBP
150 mAs 9 mGy FBP
100 mAs 6 mGy FBP
50 mAs 3 mGy FBP

## Slide 25: What else is high contrast in Abdomen?
Vessels= CTA (abdominal aorta),
Arterial phase of liver protocol CT,
CT enterography
Lower dose
Lower kV

## Slide 26: Multiphase Exam: Dose Reduction: Liver
Arterial phase:
Lower kV (100)
Liver only coverage
AEC
PV phase:
Usual or Lower kV
Longer coverage
AEC

## Slide 27: Multiphase Exam: Pancreas
Arterial phase:
Liver and pancreas only
AEC and IR technique
PV phase:
Longer coverage
AEC and IR

## Slide 28: Adrenal Protocol
Multiphase CT
Scan length:
  T11 - L2
  Adrenals only
120 kV
AEC
IR techniques

## Slide 29: CT For Hematuria: Decreasing Scan Phases
Unenhanced CT
50 mL IV contrast bolus
250 mL saline drip infusion
Wait 15 minutes (Prone)
100 mL IV contrast @ 3 mL/sec
Scan at 100-sec. delay
Maher MM, Kalra MK, et al. BJR 2004
Non-contrast
limit coverage: kidneys to SP, Lower dose (- 50%, AEC and IR)
Post Contrast:
Wider coverage
AEC

## Slide 30: Routine Abdomen: Patient factors and Dose
Size: Obese vs average or slim
  Use AEC always for mA optimization
  Lower kV in smaller patients with contrast CT helps cut dose
Arms by the side of the patient
  Keep out or Keep above the abdomen with pad not besides
Patient motion (voluntary or involuntary)
  Pitch > 1; Rotation time < 0.8 second
Contrast or non contrast scan
Prefer contrast when ok … Higher contrast helps lower dose

## Slide 31: Routine Abdomen: Use AEC
50 mAs
3.4 mGy
200 mAs
14 mGy
339 mAs
23 mGy
200 mAs
14 mGy
Average/slim
65 kg
50 mAs
3.4 mGy
Obese
115 kg
Notes: 31

## Slide 32: Lesion margin may not be as evident on low dose CT
633 mA
24 mGy
360 mA
90 mA
45 mA
14 mGy
3.4 mGy
1.7 mGy
180 mA
7 mGy

## Slide 33: Routine abdominal CT at different mAs and CTDI vol show right renal adenoma
45 mAs
22 mAs
3.4 mGy
1.7 mGy
90 mAs
120 kV
0.9:1 pitch
5 mm
FBP
24 mGy
14 mGy
7 mGy
315 mAs
180 mAs

## Slide 34: Benign Disease: Young Patients
Ex: Appendicitis
Limited coverage: L3 to SP
AEC for size adapted dose

## Slide 35: Abdominal CT Dose Management
Lower noise: Higher dose
Lesion Detection – high confidence
Higher noise = Lower dose
Rule out Abdo CT
Low contrast lesions
Use AEC to optimize
CT colonography
CT angiography
Kidney stones
Hypervascular lesions
Use AEC &/ low kV

## Slide 36: Summary
Need for clinical indication based CT protocols
Avoid routine multiphase CT protocols
Reduce scan length when possible for multiphase CT
Use AEC
Use lower kV for smaller patients particularly with iterative reconstruction techniques