## Slide 1: Radiation Dose Management in Computed Tomography
Key Aspects for Pediatric CT L10
Material of the e-learning program
IAEA, 2017
Notes: Module 10. Key aspects of pediatric CT

## Slide 2: Objectives
Why children are special when it comes to CT
Understanding key techniques for dose reduction in children
How to tailor dose to size, clinical indication, prior studies
  Chest and abdomen CT
  Head and neck CT
Notes: The objectives of this module are to understand why children are special when it comes to CT and understand key techniques for dose reduction in children.
In addition, we will learn about the process of tailoring pediatric CT protocols and radiation doses according size, clinical indication, and availability of prior studies.

## Slide 3: Why children are so special for CT?
Well…because they are just special!
Seriously: Every patient is special but children are more special … for CT
Children are at higher risk for radiation induced stochastic effect (radiation induced cancer)
  Biology: Children have more dividing cells which are more vulnerable to radiation induced effects.
  Expectancy: Children are likely to live longer than adults, so they have more chances to have cancer over their lifetime.
Notes: Although every patient regardless of his or her age is special but children are more special when it comes to CT scanning since they are at higher risk for radiation induced stochastic effect, that is radiation induced cancer.
The reason for increased risk for radiation induced cancer in children may be due to the fact that children have more dividing cells which are more vulnerable to radiation induced effects and children are likely to live longer than adults, so they have more chances to have cancer over their lifetime.

## Slide 4: Why children are so special for CT?
Children range widely in size from less than 1 kg to adult body habitus and larger
  Smaller children can be and should be scanned with smaller doses
Young children often need faster scan time or sedation or anesthesia or immobilization for CT
Gamut of pathology for CT in children are different
  Triage to other tests (radiography, ultrasound or MRI)
  Seldom need repeat scanning or multiphase CT
  Benign indications are commoner in children
Notes: Children also vary widely in size from less than 1 kg to adult body habitus and even larger. It is important to remember that smaller children can be and should be scanned with smaller doses as compared to the adult patients.
Young children often need faster scan time and/or sedation or anesthesia for CT scanning. In other words, they often need measures to reduce motion artifacts in CT scanning, which can lead to repeat scanning and missed information.
Gamut of pathology for CT in children are also different from adults. Several clinical indications in children can be triaged to safer non-radiation or low radiation based imaging tests such as radiography, ultrasound or MRI. Compared to adults, children seldom need multiphase CT protocols.

## Slide 5: Developing CT protocols in children
Children should not be scanned protocols created “on the fly”
Protocols should be pre-defined, and archived on CT
Demands from a pediatric CT are much higher
  Errors can happen
  Children can move during scanning
  Children can get anxious while waiting
  Children may need longer sedation
  Contrast injection can go wrong!
  More time will be wasted!!
  Not all radiographs and radiologists are equal in expertise
Notes: Children should not be scanned with “on the fly” protocols. All pediatric CT protocols should be pre-defined, and archived on CT scanner.
Demands from a pediatric CT are much higher and without set protocols, errors can happen, children can get anxious while waiting for scanning to begin, children may need longer sedation due to increased wait time.

## Slide 6: Best all time winner for CT dose reduction
… Good mechanism to guide appropriateness of CT!
Notes: The best way to reduce radiation dose in children is to ensure that clinical indication for performing CT is appropriate and justified and needed information can not be obtained with other safer imaging modalities in a timely, efficient and accurate manner.

## Slide 7: Specific Pediatric CT Protocols
Never…ever use adult protocols in children.
Like all adult protocols, pediatric CT protocols should also begin with clinical indications.
Each scan protocol should address:
  Clinical question to be answered
  Dose adjustment for patient size
  Specific adjustments for uncooperative children
  Special attention to contrast injection
Notes: One should never use adult protocols in children unless they are extremely large in size and lower radiation dose is likely to impair diagnostic information.
Like all adult protocols, pediatric CT protocols should also begin with clinical indications.
Each scan protocol should address
Clinical question to be answered
Dose adjustment for patient size
Specific adjustments for uncooperative children, and
Special attention to contrast injection

## Slide 8: Guiding Principles for Pediatric CT
Perform only necessary CT:
  Triage to other safer tests when appropriate
  Do not delay or hesitate in emergent valid clinical need
  Avoid unnecessary repeat studies
Scan only the organ or anatomy indicated
Properly center all children in the CT gantry
Adjust scan parameters to size, body region and indication
Use pediatric positioning accessories
Minimize multi-phase contrast CT
Notes: The most important principle for good pediatric CT practice is to perform only necessary CT and avoid unnecessary repeat examinations.
It is important to ensure that scanning is limited to the organ or anatomy indicated.
It is important to ensure proper centering of children in the CT gantry.
Scan parameters must be properly adjusted to allow dose adjustment based on size, body region and indication.
Use of pediatric positioning accessories should be encouraged to minimize motion artifacts.
In children, it is important to minimize use of multi-phase contrast CT protocols.

## Slide 9: Why Indication driven protocols?
Certain indications can be scanned at substantially reduced dose in children.
Some of these indications can be scanned at extremely low doses compared to others.
Low dose head CT | Low dose chest CT | Low dose abdomen CT
Craniosynostosis Ventricular shunt patency Paranasal sinuses Facial trauma Follow up head CT | Lung abnormalities Airway evaluation Pleural effusion/empyema Chest wall deformities Scoliosis Chest trauma Follow up chest CT | Kidney stones CT enterography Follow up abdominal CT Bony deformities
Notes: Certain indications can be assessed at substantially reduced dose in children. The table presents a list of clinical indications for which low radiation dose CT exams are sufficient. For example, in the head CT, craniostenosis, ventricular shunt patency, paranasal sinuses, facial trauma, and follow up head CT can all be scanned at lower radiation dose compared to routine head CT examinations. Evaluation of lung and airway abnormalities can also be performed at extremely low radiation dose. Likewise, indications such as kidney stone, and CT enterography can also be assessed at reduced radiation doses.

## Slide 10: Pediatric CT: Patient preparation
Address nervousness
  Be calm and professional
  Inform steps of CT to patient and accompanying person
  Do not over-crowd
  Decide if someone needs to stay in the room (when appropriate)
Minimize pain
  Take intravenous access if needed outside the room and in advance
Instruct, demonstrate and practice
  Tell what should be done (breath-hold), when appropriate
  Demonstrate how it will be done (when appropriate)
  Practice what has to be done (breath-hold, when appropriate)
Ensure contrast injection is well planned: Do not botch!!
Avoid artifacts from external and internal sources
Notes: In order to reduce motion artifacts during CT scanning in children, it is important to address nervousness
by informing steps of CT to the child and/or accompanying caretaker.
When possible, pain during CT scanning should be minimized. If possible, intravenous access if needed
should be obtained outside the room and in advance.
When appropriate in older children, instruct, demonstrate and practice the CT scanning. Inform the
patient, who they should hold their breath and avoid voluntary movements. At times, demonstration and
practice works better than mere instructions especially when it comes to breath-holding.
Ensure contrast injection is well planned and minimize risk of contrast extravasation or missed contrast
bolus, since these may require repeat scanning.
One should also avoid artifacts from external and internal sources when possible.

## Slide 11: Pediatric CT: Scanner Preparation
Choice of CT scanner
  If you can, choose the CT which is faster
    Dual source CT are faster than single source CT
    Wider area detector CT are faster than others
  If you can, choose the scanner with AEC, automatic KV selection and/or iterative reconstruction techniques
  If you can, choose an experienced or expert radiographer
  When appropriate and possible, discuss the CT protocol with a radiologist (so she/he can triage too!!)
Notes: If more than one CT scanners are available for scanning, choose the CT which is faster. For example,
scan coverage with dual source CT is often faster than the single source CT. Likewise, wider area
detector CT are faster than other scanners.
If you can, choose the scanner with automatic exposure control, automatic KV selection and/or iterative
reconstruction techniques.
If you can, choose an experienced radiographer to supervise the pediatric CT examination.
When appropriate and possible, radiographers should discuss the CT protocol with a radiologist (so s/he
can triage too!!)

## Slide 12: Pediatric CT Protocols: Good Practice
Centering is extremely important in children
Offcentering or miscentering can:
  Mess up estimation of mA for AEC
  Mess up beam shaping filters’ functions
    Asymmetric distribution of noise and artifacts
  Increase surface dose to vital organs (e.g. breasts)
  Effect can require repeat scanning especially for very low dose protocols with severe artifacts
Notes: Off-centering of the patient in the scanner is frequent problem in smaller patients. It is important to
understand that miscentering can:
Mess up estimation of tube current with automatic exposure control techniques.
Off-centering can also mess up beam shaping filters’ functions which assume good patient centering. This can lead to asymmetric distribution of noise and artifacts and increase surface radiation to vital organs (thyroid/breasts).
Effects of off-centering can require repeat scanning especially for very low dose protocols with severe artifacts.

## Slide 13: Pediatric CT Protocols: Good Practice
Indication and size specific scan parameters:
  Tube Current:
    Prefer AEC over fixed mA
    You can use fixed mA for very low dose CT protocols
    Minimized mAs for
      Craniostenosis, chest wall deformities, scoliosis
    Reduced mAs
      Follow up CT, lung/airway abnormalities, kidney stones
    Lower mAs &/or KV for iterative reconstruction than FBP
Notes: Pediatric CT protocols must allow for indication and size specific scan parameters.
With regards to tube current, automatic exposure control techniques should be used over fixed tube current. As stated in preceding slides, certain clinical indications must be scanned at low radiation doses. Iterative reconstruction techniques also enable use of lower tube current and tube potential.

## Slide 14: Pediatric CT Protocols: Good Practice
kV selection:
Automatic kV selection should be used when available
  Use appropriate settings based on clinical indication
  Use lower reference KV for its application
Manual KV selection
  Do not underuse low kV in small children
  Lowest kV should be used for very small children
    Infants can be scanned adequately at 70-80 kV
  Lowest kV should be used for CT angiography
    Most children up to 80 kg can be scanned at 80 kV
    Heavier children can be scanned at 100 kV
Iterative reconstruction enables lower kV more often
Notes: Automatic KV selection should be used when available to enable use appropriate KV settings based on clinical indication and size. Lower kV can often be used for scanners with iterative reconstruction techniques.
When using manual kV, it is important to not underuse low kV in small children. Lowest kV of 70 or 80 kV should be used for very small children.
Lowest kV should be used for CT angiography and children weighing as much as 80 kg.

## Slide 15: Pediatric CT Protocols: Good Practice
  Select non-overlapping pitch
    Non-overlapping pitch should be preferred to minimize scan duration and motion artifacts
    On certain scanners, lower pitch means higher dose
Beam collimation: Minimize scan duration
  For longer scans (chest, abdomen, spine), choose wider beam
  For variable detector array systems: Choice depends on required need for specific slice thickness
  Ex: 16 x 0.75 mm if < 1 mm slice are required
  Ex: 16 x 1.5 mm when > 1.5 mm slices required
Fast gantry rotation speed: minimize motion artifacts
Notes:   To enable faster scanning, children should be scanned at non-overlapping pitch. On certain scanners, for example GE and Toshiba, lower pitch means higher dose if other scan parameters are held constant.
  Likewise, appropriate detector configuration should be used to maximize dose efficiency and minimize scan duration.
  For longer scan ranges such as chest, abdomen, and spine, one should choose wider beam collimation.
For variable detector array systems: Choice of detector configuration should depend on need for specific slice thickness
For example in a 16-slice MDCT, a 16 *0.75 mm detector configuration is better if less than 1 mm slice thickness is required. For the same scanner 16 *1.5 mm detector may be more suitable for long scan range when 1.5 mm or thicker slices are acceptable.
A faster gantry rotation speed should be used in children to decrease scanning time and reduce motion artifacts.

## Slide 16: Pediatric CT Protocols: Good Practice
Minimize repeat scanning or multiple phase CT
    Generally are not needed in children and increase radiation dose
    Generally a single acquired series is enough
    Routine pre- & post-contrast CT should be avoided
Scan length: Must be limited to what is needed
    Chest: Lung apices to base (no need to go for adrenals)
Single run CT: if scanning two contiguous regions
    Whenever appropriate, go for a single run helical: It avoids overlapping dose, it is more dose efficient (remember over-ranging), and it is faster!
Notes: Users should minimize repeat scanning or multiple phase CT as they are not generally needed in children and their use increases radiation dose. Generally a single acquired series is enough in children and routine pre- & post-contrast CT should be avoided.
Scan length must be limited to what is absolutely necessary. In adults, routine chest extends to the adrenal glands, in children chest CT can be limited to lung bases.
When scanning two contiguous body regions in children, it is better to scan them in a single run to avoid additional radiation dose from overlapping scan series.

## Slide 17: Pediatric CT Protocols: Good Practice
Slice thickness
  Thick slices have lower noise: Use these for low contrast
  High contrast organs can be seen at greater noise
    Lungs and vessels can tolerate higher noise in thins
Reconstruction kernels
  Softer kernels have lower noise: use for abdomen
  Avoid too sharp kernels: they increase noise and may necessitate an increase in dose for future exams
Reconstruction techniques
  Use iterative reconstruction techniques when available to enable further dose reduction.
Notes: It is important to remember that thick slices have lower noise and can be used to better assess low contrast regions such as liver. On the other hand, high contrast organs such as lungs and contrast enhanced blood vessels can be seen at greater noise in thin images.
Softer reconstruction kernels have lower noise and should be used for low contrast structures such as abdomen. Routine use of excessively sharp kernels in low dose CT should be avoided as they increase noise and may necessitate one to increase dose for future exams
When available, iterative reconstruction techniques should be used to radiation dose reduction.

## Slide 18: Dose adjusted to clinical indications in chest CT
Indications | Result | Author (journal year)
Bronchiectasis | Dose reduction with low mAs | Yi et al. (AJR 2003)
Pulmonary nodule | Dose reduction with use of low mAs | Diederich et al. (Radiology 1999)
Pneumonia | Dose reduction with low kV and low mA | Rizzi et al. (Clinical imaging 2007)
Pulmonary embolism | Dose reduction with low kV and low mA | Doshi et al. (BJR 2008)
Radiation dose should be adjusted according to the clinical indications.
Notes: A number of chest CT indications as illustrated in the table can be assessed at substantially lower radiation doses such as bronchiectasis, pulmonary nodules, pneumonia, and pulmonary embolism. This implies that indication based scan protocols can help reduce dose substantially.

## Slide 19: Dose adapted to the clinical indications for abdomen
Indications | Results | Authors (journal year)
Appendicitis | Tube current can be reduced substantially to achieve lower dose | Fefferman et al. (Radiology 2005) Keyzer et al. ( Radiology 2004)
Colon polyps | Use of lower mA and AEC help reduce dose | Graser et al. (AJR 2006)
Renal calculi | Use of lower mAs and AEC help reduce dose versus routine abdomen | Kalra et al. (Radiology 2005) Heneghan et al. (Radiology 2003)
Radiation dose should be adapted according to the clinical indications.
Notes: A number of abdominal CT indications as illustrated in the table can be assessed at substantially lower radiation doses such as urinary calculi, colonic polyps on CT colonography, and appendicitis. This implies that indication based scan protocols can help reduce dose substantially. A note of caution regarding appendicitis evaluation of CT is important. It is important to first use ultrasound to assess presence of appendicitis. Failure to visualize appendix can trigger MR as the next line imaging and CT if MR is not possible or available. When performing CT for appendix evaluation, it is important to ensure limited scan coverage to reduce radiation dose.

## Slide 20: Dose adapted to number of prior CT scans
Follow up of : | Results | Authors (journal year)
Cystic fibrosis | Low mAs and skipped acquisition (axial scanning) | Jimenez et al. (Arch. dis. Child. 2006)
Ventriculoperitoneal shunt | Low dose from use of low mAs and KV | Udayashankar et al. (AJNR 2008)
Craniosynostosis | Use of low kV and low mA | Trattner et al. JACR 2014
Notes: For clinical indications requiring follow up imaging such as cystic fibrosis, ventricular shunt patency, craniosynostosis, and inflammatory bowel diseases in children, it is important to reduce radiation dose with follow up CT scanning.

## Slide 21: Size based CT dose optimization in Children
Frush et al. described use of Broselow-Luten System* for color coded system for optimizing radiation dose in children in a seminal paper in AJR 2002.
The system outlined an outstanding way of using height and weight of children to optimize kV and mA.
This system relates the height of the child to the weight to provide important information on drug doses, and most suitable equipment sizes.
Frush et al. AJR 2002; 178: 721-726
Notes: Dr. Frush and colleagues were the first to describe a color coded system for optimizing CT protocols and radiation doses in children using Broselow-Luten System. The system outlined an outstanding way of using height and weight of children to optimize tube current and tube potential for pediatric CT protocols.

## Slide 22: Color coded dose optimization in children
Other authors have described color coded system in children to adapt radiation dose to
  Patient weight
  Body region
  Clinical indication
  Number of prior exams
  Automatic exposure control (AEC)
Singh et al. Radiology 2009
Singh et al. Radiology 2012
Notes: Other authors have described color coded system in children to adapt radiation dose to
Patient weight,
Body region,
Clinical indication, and
Number of prior exams, while using automatic exposure control technique.

## Slide 23: Color coding of indications, body region & prior CT
Indications | Zone number | Color zone | Radiation dose
Routine or “rule out” abdomen situations | 1 | Pink | Standard
Follow up abdomen CT, first chest CT | 2 | Green | 
Multiple prior chest & abdomen CT scans, Index CT for bone indications | 3 | Red | 
CT angiography | 4 | Grey | 
Color codes help remember the different zones based on body regions, clinical indication & number of prior CT.
Singh et al. Radiology 2009
Singh et al. Radiology 2012
Notes: In this color coded indication, body region, and prior CT based protocoling system, CT protocols were initially binned into color zones including pink (for initial abdominal CT), green (for follow up abdominal CT or initial chest CT), and red (for patients with multiple prior CT and for evaluation of bony deformities) zones, in the order of decreasing radiation doses.

## Slide 24: Split each indication based protocol into different weight groups
Weight (kg) | Subcodes / Subzones
< 9 | Babies
10-26 | Cuties
27-45 | Kiddies
46-100 | Teenies
>101 | Biggies
zone / numeric code
Pink /1
While some AEC techniques do not require specific splitting of indication groups into weight groups, each weight group also had different kV to optimize dose.
Notes: After receiving the information on the specified color zone, radiographers then weighed the child to determine the appropriate weight category protocol for each color zone. Although automatic exposure control technique was used, individual settings for each weight group was important to ensure that an optimal combination of kV and tube current is used for each child.

## Slide 25: Adjust parameters to patient size for each color zone
Weight (kg) | Pink zone | Noise index | Min- max mA | kV
< 9 | Babies | 5 | 65-130 | 80
10-26 | Cuties | 7 | 80-160 | 100
27-45 | Kiddies | 10 | 95-190 | 120
46-100 | Teenies | 12 | 110-220 | 120
>101 | Biggies | 15 | 125-300 | 120
Weight (kg) | Green zone | Noise index | Min- max mA | kV
< 9 | Babies | 7 | 50-100 | 80
10-26 | Cuties | 9 | 60-120 | 100
27-45 | Kiddies | 11 | 70-140 | 120
46-100 | Teenies | 13 | 80-160 | 120
>101 | Biggies | 16 | 90-240 | 120
Routine abdomen indications
Follow up abdomen
First chest CT
While some AEC techniques do not require specific splitting of indication groups into weight groups, each weight group also had different kV to optimize dose.
Notes: The above mentioned tables present two zones, the pink and green zones reported with this system of color coding pediatric CT. In each zone, there were subzones for different weight groups. Each weight sub-group used different settings of automatic exposure control technique albeit with extremely conservative use of lower tube current.

## Slide 26: Classification of indication category: Radiologist’s job!
Offers opportunity to triage patient to other modality
Radiologist should look at CT requests for
  Clinical indication
  Presence of prior CT
  Body region
Based on indications need for CT should be decided:
CT not needed
CT needed
No CT, No additional dose
Call physician to triage appropriately
  Specify
Notes: This color coding system required that the radiologists review the clinical indication for each pediatric CT and specify a color zone to the radiographer. This also gave the radiologists an opportunity to assess appropriateness of clinical indication for CT as well as possible triaging opportunity to other imaging modalities.

## Slide 27: Raising awareness amongst radiographers
Imbibe need for dose reduction specially in children
Understand the color zones specified by the radiologists
Weigh all children, avoid guessing
Only guess, if weight cannot be measured
Pick exact scan length, do not under or over “scan”
Center patients in CT for good dose reduction with AEC
Select radiologist specified protocol & weight determined subzone
Do not make any changes to parameters unless specified
Notes: Such clinical indications driven clinical protocols in children imbibe need for dose reduction specially in children. For the radiographer, it required them to weigh all children, and limit guessing only when weight could not be measured.
The pediatric CT protocols should also convey to the radiographer the need to pick exact scan length, do not under or over “scan,” appropriately center children in CT for good dose reduction with automatic exposure control, select radiologist specified indication based color zone and weight-based specific CT protocol.

## Slide 28: Save protocols on the scanner (lock down)
Color zone
Pink
Green
Red
Yellow
Blue
Grey
Weight subzone
<9 kg
10-26 kg
27-45 kg
46-100 kg
>101 kg
Weight (kg) | Green zone | Noise index | Min- max mA | kV
< 9 | Babies | 7 | 50-100 | 80
10-26 | Cuties | 9 | 60-120 | 100
27-45 | Kiddies | 11 | 70-140 | 120
46-100 | Teenies | 13 | 80-160 | 120
>101 | Biggies | 16 | 90-240 | 120
CT scanner user interface
Notes: When creating a sophisticated system for CT protocols for children, it is important to appropriately name the CT protocols and archive them on the scanner user interface as shown in the slide. Some scanners allow users to lock down the protocols on the scanner with password protection.

## Slide 29: CT Workflow
Some, Child CT
Electronic Medical Records
Indications
Number of prior CT exams
We need CT!
Red Zone!
Oh, Baby!
1
2
3
4
Weigh or guess
Notes: In summary, given the increased risk from use of CT in children, referring physicians should use of evidence based guidelines for ordering CT in children. Radiologists should review order for pediatric CT for appropriateness of clinical indication and presence of prior imaging. If CT is deemed necessary, radiographers should be guided on use of suitable indication based clinical protocol. Radiographer should then select appropriate weight based protocol archived on the scanner user interface.

## Slide 30: Attempt gradual dose reduction in small steps
Pink zone | Weight (kg) | Noise index | Min - Max mA | kV
Babies | < 9 | 4 | 65-130 | 80
Cuties | 10-26 | 6 | 80-160 | 100
Kiddies | 27-45 | 8 | 95-190 | 120
Teenies | 46-100 | 10 | 110-220 | 120
Pink zone | Weight (kg) | Noise index | Min - Max mA | kV
Babies | < 9 | 5 | 65-130 | 80
Cuties | 10-26 | 7 | 80-160 | 100
Kiddies | 27-45 | 10 | 95-190 | 120
Teenies | 46-100 | 12 | 110-220 | 120
Example: In MGH Boston, doses for each indication was reduced in 3 steps over a period of one year
Pink zone | Weight (kg) | Noise index | Min- max mA | kV
Babies | < 9 | 5 | 65-130 | 80
Cuties | 10-26 | 7 | 80-160 | 100
Kiddies | 27-45 | 10 | 95-190 | 120
Teenies | 46-100 | 12 | 110-220 | 120
Biggies | >101 | 15 | 125-300 | 120
Step 1
Step 2
Step 3
Notes: Upon initiation of dose and protocol optimization of pediatric CT, it is important to start conservatively and then decrease the radiation dose in a stepwise manner. The slide illustrates the three step radiation dose reduction in children over a period of one year.

## Slide 31: Assess dose reduction following implementation of CT protocol
CTDI vol (mGy) | Dose reduction
6.7 | 61%
17.4 | 
Following implementation of CT protocol:
Compare radiation dose with old protocols, compliant & non compliant with composite color coded protocols.
 | CTDI vol (mGy) | Dose reduction
Compliant | 6.6 | 65%
Non compliant | 19.1 | 
* Results of color coded protocol implemented at MGH
Step 2
Step 3
Step 1
Singh et al. Radiology 2009
Singh et al. Radiology 2012
CTDI vol (mGy) | Dose reduction
4.6 | 75%
18.5 | 
Notes: The slid e illustrates the extent of dose reduction at each step. It is important to monitor pediatric CT doses to ensure that they are in compliance with the CT protocols. In the above slide, children scanned without compliance with the pediatric CT protocols received substantially higher radiation dose compared to those scanned in compliance with pediatric CT protocols.

## Slide 32: Pink zone abdominal CT
13 year old girl weighing 70 kg (transverse diameter 33.2 cm) underwent abdominal CT for abdominal pain
120 kV 282-485 mA
Noise index: 28
CTDIvol: 16.5 mGy
120 kV 183-219 mA
Noise index: 12
CTDIvol: 10.8 mGy
Non compliant
Pink zone
Notes: Radiation dose in a child scanned with adult CT protocol are significantly greater than needed. In this example, pelvic abscesses are well seen in lower radiation dose CT performed with pediatric protocol (pink zone) as compared to higher dose adult CT protocol on the left side.

## Slide 33: Green zone chest CT
16 year old boy weighing 76 kg (transverse diameter 31.4 cm) underwent chest CT for follow up of pulmonary nodule
140 kV 439 mA
Noise index: 12
CTDIvol: 29.3 mGy
120 kV 160 mA
Noise index: 10
CTDIvol: 5.4 mGy
Non compliant
Green zone
Notes: Radiation dose in another child scanned with adult CT protocol are significantly greater than needed. In this example, lung nodule in the left lower lobe is well seen in significantly lower radiation dose CT performed with pediatric protocol (green zone) as compared to higher dose adult CT non-compliant protocol on the left side.

## Slide 34: Red zone abdomen CT
120 kV, 93-119 mA
Noise index : 9
CTDIvol : 3.8 mGy
120 kV, 45-62 mA
Noise index: 12
CTDIvol : 1.8 mGy
Red zone
Green zone
2 year old boy with transverse diameter 17.8 cm underwent abdominal CT for post left adrenal neuroblastoma resection
Notes: This patient underwent serial follow up abdominal CT for assessing local recurrence or metastatic cancer following resection of left adrenal neuroblastoma. The image on the left presents the first follow up protocol in green zone at 3.8 mGy. The right side image has higher noise but shows the region of interest equally well with red zone protocol at CTDI vol of 1.8 mGy. Patients undergoing follow up CT should be scanned at lower radiation dose when possible.

## Slide 35: Red zone chest CT
100 kV, 80 mAs
Noise index : 7
CTDIvol : 4.7 mGy
100 kV , 50 mAs
Noise index: 8
CTDIvol : 2 mGy
Red zone
5 year old boy with transverse diameter 17.4 cm underwent chest CT for follow up of lung nodule
Green zone
Notes: Five year old boy underwent chest CT for assessing lung nodules. The image on the left presents the first chest CT protocol in green zone at 4.7 mGy. The right side image acquired with red zone protocol at CTDI vol of 2 mGy also demonstrates the right apical nodule equally well. Patients undergoing follow up CT should be scanned at lower radiation dose when possible.

## Slide 36: Pedi Neuro CT
Setty, kalra, et al. SPR 2008
Trauma
VP shunt patency
Post operative
Notes: When MR is available for use and not contraindicated, routine head CT in children should be limited to trauma, evaluation of ventricular shunt patency or post operative situations. Evaluation of ventricular size following placement of ventricular shunts can be perfomed at very low dose as well. Both low fixed tube current or automatic exposure control technique can be used to reduce radiation dose for such exams. The left image for routine head CT was acquired with fixed tube current with resultant CTDIvol of 33 mGy. The right image was acquired with automatic exposure control at CTDI vol of 13 mGy with shunt patency evaluation protocol. Recent studies have also reported an extremely fast head MRI protocol for evaluation of shunt patency in a matter of few minutes.

## Slide 37: Pedi skull
Notes: Indications for evaluation of bony abnormalities such as pectus excavatum and craniosynostosis can be scanned at extremely low radiation dose levels. In this slide, an infant underwent head CT at 50 mAs and 80 kV with DLP of 20 mGy.cm for the entire scan. Although resultant images do not provide sufficient image quality for evaluation of brain parenchyma, bony details are sufficient to assess absence of suture fusion in this patient.

## Slide 38: Chest CT for aorta & pulmonary arteries
13 month old girl with pulmonary artery stenosis scanned at 80 kV, low mA (AEC) and without ECG gating
Notes: A 13 month old girl underwent chest CT at 80 kV with automatic exposure control technique. The low dose images have high amount of noise but are sufficient to detect enlarged right atrium and ventricle with stenosis of the left pulmonary artery at its origin from the main pulmonary trunk.

## Slide 39: Cardiac CT with high pitch and low kV
Age 14 years Weight 59 kg EKG gated heart Pitch 3.4 80 kV CTDIvol = 2.5 mGy | Age 65 years Weight 135 kg Non gated pulm. veins Pitch 3 100 kV Dose 3 mSv
Notes: CT coronary angiogram in a 14 year old boy on a dual source multidetector row CT at 80 kV and high non-overlapping pitch of 3.4. Low dose examination was sufficient to rule out anomalous coronary arteries as shown in the bottom left image.
The images on the right demonstrate use of similar technique but at higher kV of 100 in an elderly man weighing 145 kg. Evaluation of pulmonary veins is sufficient at low radiation dose despite large patient size.

## Slide 40: Summary
Pediatric CT should be performed once it is clear that other safer modalities can not give accurate and timely information.
Adult CT protocols and radiation doses should not be used for pediatric CT.
Use of multiphase CT must be avoided in children.
Iterative reconstruction, AEC, and automatic kV selection techniques must be used in children when available.
Notes: In summary, pediatric CT should be only performed once it is clear that other safer modalities can not give accurate and timely information.
Adult CT protocols and radiation doses should not be used for pediatric CT
Routine use of multiphase CT must be avoided in children.
Iterative reconstruction, automatic exposure control, and automatic kV selection techniques must be used in children to reduce radiation dose when they are available.