## Slide 1: Radiation Dose Management in Computed Tomography
L07
Chest CT: Key Aspects
Material of the e-learning program
IAEA, 2017

## Slide 2: Objectives
Why is chest so special for CT?
Understand key protocols for chest CT
Understand key scan parameters for optimizing radiation dose for chest CT:
  Patient size based dose optimization
  Clinical indication based protocol development

## Slide 3: What is so special about Chest CT?
Lungs:
  Air causes less x-ray attenuation and so requires less radiation dose
  Air results in high inherent tissue contrast versus most lesions and so can be imaged at lower dose
Heart, large blood vessels and lungs:
  Pulsating structures require faster scanning
  Short scan duration reduce respiratory motion artifacts as well
  High contrast offers opportunities for low dose

## Slide 4: Chest CT doses must be different from Abdomen
Routine chest CT dose < routine abdomen CT doses
Low dose Lung nodule CT < low dose kidney stone CT
Chest CT angiography < Abdomen CT angiography
Chest should have fewer multi-phase or pass studies

## Slide 5: Developing CT protocols
“On the fly” protocol creation is not optimal
Valuable time can be wasted for urgent CT
Level of training between radiographers and between radiologists is different
Errors can result in implementation since there are far too many scan parameters
Consistency is difficult to achieve:
  Same clinical indication and size: Similar doses
  Different clinical indications or size: Different doses

## Slide 6: All time winner for The BEST CT dose reduction strategy……..
Appropriate clinical indication for CT

## Slide 7: Specific Chest CT Protocols
All scan protocols must begin with the clinical indications for their use.
Scan protocols must be based on clinical indications
Each scan protocol should address
  Number of scan phases required
  Scan range for each phase
  Scan parameters for each phase
  Dose adjustment based on patient size

## Slide 8: Need: Indication driven protocols?
Certain things are ok at lower dose (lung nodules)
Others need higher dose (mediastinal LN)
Radiation doses and chest CT protocols:
CT Lung nodule follow up < Routine Chest CT post contrast < CT Pulmonary embolism < Routine chest CT Non Contrast

## Slide 9: Chest CT protocols: Good Points
Indication and size specific scan parameters
Tube Current:
  Prefer AEC over fixed mA
  Can use fixed mA for very low dose CT protocols
  Adjust mA for different clinical indications
    Lower mAs for nodule follow up/lung cancer screening
    Lower mAs for iterative reconstructions than FBP
kV : Manual selection or Automatic KV selection
  Use lower kV for contrast enhanced chest CT
  Use lower kV for CT angiography unless patient is large

## Slide 10: Chest CT protocols: Good Points
Pitch for chest CT
Non-overlapping pitch should be preferred to minimize scan duration and motion artifacts
    Generally pitch for chest CT should be ≥ 1:1
Number of scan series for chest CT
    Generally a single acquired series is enough for chest CT
    Don’t acquire routine pre-contrast prior to post-contrast
    Diffuse lung Ds: Dose should not be multiple folds higher
Scan length: Restricted to what is needed
    Routine chest: Lung apex to Adrenal glands
    Lung nodule follow up, Pulmonary embolism, and lung cancer screening: Lung Bases to Apex

## Slide 11: Chest CT protocols: Good Points
Beam collimation:
  Wider beam collimation for most routine CT: More efficient (esp. < 64 slice)
  For variable detector array systems: Choice depends on need for specific slice thickness
    Ex: 16 * 0.75 mm if < 1 mm slice are required
    Ex: 16 * 1.5 mm when > 1.5 mm slices required
Gantry rotation speed:
  Fast to minimize motion artifacts (0.4 - 0.5 s)
Reconstruction kernels:
  Standard or detail: thinner slices (cardiac CT or CTA) or lower dose
  Sharper: Bones and lungs

## Slide 12: Indication Driven CT Protocols
Protocol | Clinical Reasons | Specific instructions
Routine chest with IV contrast | Masses, infections, trauma to lungs, mediastinum, & pleura | No pre-contrast images
Routine chest without contrast | Elevated creatinine for above, Follow-up nodule in pt’s with CA | No post-contrast images
Lung nodule follow up | Follow up nodule without known malignancy | Do not cover adrenals
Diffuse lung disease protocol | For diffuse lung diseases e.g. Sarcoid, bronchiolitis obliterans, ILD, pulmonary fibrosis, | Helical: inspiration Axial: Expiratory Axial: Prone images
Pulmonary Embolism | Suspected or known pulmonary embolism | Do not cover adrenals Scan from lung bases
Tracheal protocol | Tracheobronchomalacia Tracheal stenosis | Lower dose Inspiratory & expiratory images

## Slide 13: Know thy scanner
AEC
KV settings and techniques
Iterative reconstruction
Other tips and techniques

## Slide 14: Search for the Optimal Blend: The Magic Protocols?
Parameter | Value
Scanner | Magique Optimale
Profile | 1024 Stealth mode
KV | Scanner determined
mA | Scanner determined
Rotation time | Scanner determined
Slice thickness | 0.1 mm, and 5 mm
Reconstruction kernel | Blend Edge sharp, substance soft
Recon Technique | MRT (magical recon technique)
CTDI vol | 0.1 mGy
DLP | 6.0 mGy.cm

## Slide 15: Localizer Radiograph
Remember good “centering” = good AEC and quality
Reduce dose for localizer radiograph
  Two localizers are better than one for good centering
  Considerable variation between how different vendors use these for AEC: Some use both localizers and others use the last one

## Slide 16: Dose Reduction with Axial Mode
After lesion localization, reduce dose for CT guided Bx
  Axial acquisitions
  Reduce scan length and mA &/or kV
Initial helical to localize lesion
Bx needle localization with axial mode
Axial
100 kV
40 mAs
Helical
120 kV
250 mAs

## Slide 17: Good CT: Diffuse lung disease Helical vs Axial Mode
Scan parameters | Values
Scan coverage | Apices to adrenals
Mode | Helical
Time | 0.5 second
Recon. thickness | 2.5 mm
Detector collimation | 64*0.625 mm
Pitch | 0.984:1
Speed | 40 mm/rotation
KVp | 120
Recon. kernel | FBP or h-IRT
Patient Weight | AEC settings
<60 kg | 32 NI (100-200)
61-90 kg | 35 NI (100-250)
>91 kg | 38 NI (100-400)
Helical: Most chest CT
Axial: Diffuse lung Dz
Prospective EKG triggering
CT guided biopsy
Inspiration Helical
DLP= 419 mGy.cm
Expiration AXIAL
DLP= 86 mGy.cm
Prone Axial
DLP= 43 mGy.cm
MGH Boston

## Slide 18: Scan Length
Apex to adrenals:
Routine Chest
Apex to lung base:
Pulmonary Embolism
Lung nodules
Lung Cancer screening
Benign lung disease

## Slide 19: Limiting Scan Length
  Shorter length = Smaller Dose (when all other scan parameters are constant).
Routine Chest CT
Apices to adrenal
PE protocol
Apices to lung bases

## Slide 20: Scan Overlap
Namasivayam et al, ARRS 2006
Neck-Chest CT
Chest-Abdomen CT
Minimize scan overlap to reduce dose!
Notes: 20

## Slide 21: Tube Current for Chest CT
Substantial mA reduction:
Pulmonary nodule follow up
Pediatric lungs
Follow up cystic fibrosis
Pleural effusion drainage
Lung cancer screening
Central airways
Needle localization for biopsy
Controlled mA reduction:
Mediastinal abnormalities
Lymph node evaluation
Large patient
Cancer staging

## Slide 22
AEC versus Fixed Tube Current
Most chest CT should be performed with AEC
  AEC optimizes dose to patient size
  You optimize AEC to clinical indication
Know the quirks, ins and outs of your AEC
  Lower image quality is needed for chest CT than for abdomen
  Lower image quality is needed for lung nodule than routine chest
Some indications, like lung nodule follow up CT can be scanned with very low fixed mAs

## Slide 23: Raise the Arms for AEC
20 mGy
14 mGy
20 sec later with
no change in scan factors
Kalra et al. AJR 2004

## Slide 24: Lung cancer screening CT
Scan parameters | General | Values (Ex)
Scan coverage | Apex to base | Apex to lung base
Mode | Helical | Helical
Time | Faster the better | 0.5 second
Pitch | Pitch ≥ 1 | 0.984:1
Speed | Faster is better | 40 mm/rotation
Recon. thickness | 1 and 3 mm | 1 or 1.25mm and 3 mm
Detector collimation | Wider is more efficient | 16 X 1.25mm
KV | 100-120 | 120
AEC or fixed mA | Either | AEC better than fixed mA
Target CTDI vol | 1.5 - 3 mGy | -
Substantially lower mAs compared to routine chest CT
With iterative reconstruction, 30-40% lower dose possible

## Slide 25: Lung nodule assessment on CT
120 kV/ 100mAs /6.2mGy
Mixed solid and cystic nodule in the left lower lobe abutting major fissure (circle)
Groundglass nodular lesion in the left lower lobe (arrow)
120 kV/ 75mAs/ 4.2mGy
120kV/ 40mAs/2.1 mGy

## Slide 26: Lung nodule follow up CT
Scan parameters | General | Values (Ex)
Scan coverage | Apex to bases | Apex to lung bases
Mode | Helical | Helical
Time | Faster the better | 0.5 second
Pitch | Pitch ≥ 1 | 0.984:1
Recon. thickness | 1 and 3 mm | 1 or 1.25mm and 3 mm
Detector collimation | Wider is more efficient | Wider is better if at least 1.25 mm thin slices
KV | 100-120 kV | 120
AEC or fixed mA | Either | Prefer AEC over fixed
Target CTDI vol | 1.5 - 3 mGy (FBP) Target lower for iterative reconstruction | -
Follow up protocol should be well defined
Whether Fleishner or ELCAP for > 4mm nodules

## Slide 27: Low Dose for Lung Nodules
Low fixed mAs may be as good as AEC for lung nodules
To optimize radiation dose, also have good follow up guidelines for lung nodules
Kubo et al. AJR 2008
Study & year | kV | Standard mAs | Smallest aceptable mAs for lung nodules
Rusinek (1998) | 120 | 200 | 20
Gartehnschlanger (1998) | 120 | 200 | 30
Diederich (1999) | 120 | 100 | 25
Nitta (1999) | 120 | 50 | 6
Itoh (2000) | 120 | - | 20
Karabulut (2002) | 120 | 200 | 50

## Slide 28: Lung Findings at Low Radiation Dose
150 mAs (12 mGy)
110 mAs (9 mGy)
75 mAs (6 mGy)
40 mAs (3 mGy)
All images: FBP, 2.5mm, 120 kV, 0.9: 1 pitch

## Slide 29: Lung Findings at Low Radiation Dose
120kV/110 mAs/8 mGy
120 kV/ 30 mAs/2 mGy
All FBP images
3 mm slices
B31f kernel
120 kV/75mAs/5 mGy

## Slide 30: Mediastinum at Low Dose
FBP
mA 120 kV
CTDI vol = 1.4 mGy
Esophageal mass
Pericardial and pleural effusions

## Slide 31: Mediastinum
FBP 540 mA 120 kV
CTDI vol = 20 mGy
FBP 270 mA 120 kV
CTDI vol = 10 mGy
FBP 135 mA 120 kV
CTDI vol = 5 mGy
FBP 72 mA 120 kV
CTDI vol = 2.7 mGy
FBP 36 mA 120 kV
CTDI vol = 1.4 mGy

## Slide 32: Routine chest CT
120kV/114 mAs/8 mGy
120 kV/ 40 mAs/2.8 mGy
120 kV/87 mAs/6 mGy

## Slide 33: kV Reduction: Chest CT
No use for 140 kV for chest CT : Kills contrast
Routine chest CT in children: 80 or 100 kV (70 kV on some CT)
Routine chest CT with contrast in adults ≤100 kV for most patients
Pulmonary embolism CT: ≤100 kV for <80 kg
140 KV: Smooth but nothing extra!
120 KV: Mottled ok but nothing lost!
35% lower dose

## Slide 34: Automatic kV selection
    Care kV, Siemens and KV Assist, GE
    System automatically identifies optimum kV setting (Care kV)
      Body habitus (from planning radiograph)
      Exam type (non-contrast, bone, standard contrast, vascular)
      System limits (tube current)
    System automatically proposes kV and corresponding mAs values
Notes: 34

## Slide 35: Tube potential (kV): Chest CT
Chest CT KV
Children (esp.<50 kg) 80
Adults 80 (< 50 kg); 100 (50-80 kg) 120 (> 80 kg or > 30 BMI)
Almost never 140
Encourage use of lower kV : Lung nodules FU
  Post contrast chest CT
  Chest CT angiography
  Pediatric chest CT

## Slide 36: Detector Configuration and Beam Collimation
Prefer wider beam collimation
Chest: Wider 64 >32 > 16
On 16 slice CT or lower, choice depends on desired slice width
<1mm: smaller width- less dose efficient
>1mm: wider width- more dose efficient
Scan parameters | Values
Scan coverage | Apex to adrenals
Mode | Helical
Time | 0.5 second
Recon. thickness | 2.5 mm
Detector collimation | 64*0.625 mm
KV | kV Assist
Recon. kernel | FBP or IRT
Patient Weight | AEC settings
<60 kg | 32 NI (100-200)
61-90 kg | 35 NI (100-250)
>91 kg | 38 NI (100-400)

## Slide 37: Pitch
Scan parameters | Values
Scan coverage | Apices to adrenals
Mode | Helical
Recon. thickness | 2.5 mm
Pitch | 0.984:1
Speed | 40 mm/rotation
KV | kV Assist
Recon. kernel | FBP or IRT
Patient Weight | AEC settings
<60 kg | 32 NI (100-200)
61-90 kg | 35 NI (100-250)
>91 kg | 38 NI (100-400)
Many but not all scanners adapt mA to keep dose constant irrespective of pitch
Scan speed and not dose should govern pitch
  Pitch close to 1 suffices
EXCEPTIONS:
Large patient: lower pitch
Highest pitch and fastest rotation time for uncooperative patients and children are optimal

## Slide 38: Scan Phases and Section Thickness
Scan series : Generally 1 series enough
  No routine non-contrast CT before post contrast images
  Diffuse lung Diseases: Helical (low mA) or Axial (1 @ 10-20mm)
  Tracheal protocol: Use lower dose overall, esp. expiration
Slice thickness: Acquire thin, reconstruct thick and MPR
  Beware of some scanners where prospective slice thickness is linked to mAs with AEC (Auto mA, GE)

## Slide 39: Chest CT reconstructed at various slice thicknesses
5 mm
3 mm
0.75 mm
3 mm
5 mm
Acquire thin images but read off the thick images to reduce noise in low dose images.
Thinner sections: Details & sharpness
Thicker sections: Lower noise
150 mAs, 10 mGy
40 mAs 2.7 mGy
0.75 mm
Notes: 39

## Slide 40: Dose and Reconstruction Algorithms
FBP reconstruction algorithms have limited dose reduction capabilities due to
  Higher image noise in lower dose images
  Lower artifact suppression
Newer Iterative Reconstruction (IR) algorithm
  Lowers image noise as well as artifacts
  At least 30-50% dose reduction with IR versus FBP technique
  Start with low strength of IR at the outset and then gradually increase the strength and dose reduction.
  Most modern MDCT scanners now come with 1 or more IR techniques

## Slide 41: Chest CT with FBP and iterative reconstruction
150 mAs (12 mGy)
110 mAs
75 mAs
40 mAs (3 mGy)
FBP
ASiR 50%
Available IRT from one vendor (ASIR, GE Healthcare)
150 mAs (12 mGy)
110 mAs
75 mAs
40 mAs (3 mGy)

## Slide 42: Chest CT with FBP and iterative reconstruction
FBP 150 mAs
FBP 75 mAs
FBP 40 mAs
IRIS 150 mAs
IRIS 75 mAs
IRIS 40 mAs
10 mGy
5 mGy
2.5 mGy

## Slide 43: Chest CT with FBP and iterative reconstruction
100 mA FBP 8 mGy
40 mA FBP 3.5 mGy
Safire S1
40 mA 3.5 mGy
Safire S4
40 mA 3.5 mGy
Safire S3
40 mA 3.5 mGy
Safire S2
40 mA 3.5 mGy

## Slide 44: Chest CT with FBP and iterative reconstruction
Lungs look great at 0.4 mGy in 80 kg patient
FBP 4mGy
AIDR3D STD 4mGy
FBP 0.4 mGy
AIDR3D STR 0.4 mGy
FBP 0.9 mGy
AIDR3D STR 0.9 mGy

## Slide 45: Summary: Chest CT dose reduction
Chest CT can be done at substantially lower dose
Appropriate CT indication comes first for dose reduction
Appropriate scan protocols come second
Lower dose chest CT is possible on all CT systems
Newer techniques will help cut the dose even further