## Slide 1: Radiation Dose Management in Computed Tomography
L05
Detector Configuration, Pitch and Speed in CT
Material of the e-learning program
IAEA, 2017

## Slide 2: Objective
To understand meaning and effects of:
  Detector Configuration
  Pitch
  Table feed

## Slide 3: Detector Configuration
Detector configuration of multidetector-row helical CT can be defined as the number of data channels along patient length (or z-axis) multiplied by the effective detector row thickness of individual data channel.
Represented as
  Number of data channels x effective detector row width
Determines the x-ray beam width or beam collimation
GE | Detector configuration
Hitachi | Detector configuration
Neusoft | Collimation (N x T)
Philips | Collimation (N x T)
Samsung Neurologica | Detector configuration
Siemens | Detector configuration
Toshiba | Detector configuration

## Slide 4: Detector Configuration
In variable detector arrays MDCT (scanners with variable effective detector row width), wide detector configuration can limit the z-axis resolution (minimum section thickness)
  Ex: For a 16-channel MDCT scanner,16 x 1.25 mm (20 mm) is wider and more dose efficient than 16 x 0.625 mm (10 mm) but 20 mm detector configuration will not provide <1.25mm slices!
  Thus, decision regarding detector configuration should be based on required slice thickness for such scanners.

## Slide 5: Table Feed
Table feed in helical scan mode is defined as distance travelled by the gantry table in a 360° rotation of the x- ray tube
Measured as mm/rotation
Also called table speed or couch feed
GE | Speed
Hitachi | Table speed
Neusoft | Table speed
Philips | Table speed
Samsung Neurologica | Table speed
Siemens | Table feed
Toshiba | Couch speed

## Slide 6: Pitch
Pitch for helical CT is defined as a ratio of table feed per 360° rotation of the x- ray tube and total nominal beam width
Pitch = Table feed (mm)/nominal beam width (mm)
Pitch has no units
  Changes in pitch should be made based on requirement of scanning speed rather than dose!
GE | Pitch
Hitachi | Pitch
Neusoft | Pitch
Philips | Pitch
Samsung Neurologica | Pitch
Siemens | Pitch
Toshiba | CT pitch factor

## Slide 7: Pitch: Special considerations
On single x-ray tube MDCT, pitch is generally up to 1.5:1, since greater pitch will lead to skipped data
On dual source MDCT, pitch> 1.5:1 is feasible, since two independent helices fill the gaps in data
  Such high pitch (>1.5:1) leads to much faster scanning at substantially lower dose compared to lower pitch values
    Cardiac CT angiography
    Scanning unco-operative patients
    Scanning little children without need for sedation

## Slide 8: Pitch: Caveats
Some vendors (Philips and Siemens) completely compensate for changes in pitch, so dose does not change with pitch change
  Change pitch to make the scan faster or slower
Other vendors (GE and Toshiba) do not compensate entirely for change in pitch, so higher pitch has lower dose and vice versa
  Change in pitch results in change in dose and speed

## Slide 9: Detector configuration, Table feed and Pitch
Inter-linked parameters on MDCT scanners
Implies that change in one parameter often results in change in other parameters
1.5 pitch
7 mGy
0.75 pitch
10 mGy
30% less dose

## Slide 10: Effects on Speed and Dose
Parameters | Change | Radiation dose | Scanning speed
Pitch | + - | No change for some Decreased dose for some No change for some Increased dose for some | Increased Decreased
Det. Configuration | + - | Greater dose efficiency Greater over-ranging Lower dose efficiency Lower over-ranging | Increased Decreased
Table feed | + - | No direct effect on dose | Increased Decreased

## Slide 11: Over - Ranging
Over-ranging may be defined as additional radiation dose due to extra half rotation at the beginning and the end of a helical CT acquisition
The additional half rotations are required to generate the first and last slices of helical acquisition
The half rotations results in patient dose without generation of images
Schilham et al. Radiographics 2010

## Slide 12: Over-Ranging
Extent of over-ranging is directly proportional to pitch and beam width
  Higher pitch results in greater over-ranging
  Wider beam width has greater over-ranging
 | Direction in scan range | Organs affected by over-ranging
Head CT | Above head Below head | None Eyes
Chest CT | Above chest Below chest | Thyroid Abdominal organs
Abdomen CT | Above abdomen Below abdomen | Breasts Testes in males
Schilham et al. Radiographics 2010

## Slide 13: Over-ranging: Implications on scan parameter choice
For short scan length: over-ranging can represent a significant portion of overall dose
  Thus, short scan ranges should be scanned at lower pitch and lower beam collimation
    Ex: Small joints, small regions of interest, head CT
For longer scan length: over-ranging represents a small portion of overall dose
  Thus long scan ranges can be scanned with faster pitch and wider beam collimation to decrease scan times.
    Ex: Chest CT, Abdomen CT, extremity CT

## Slide 14: Over-ranging: Solutions
Adaptive dose shields (Siemens and Philips):
Dynamic beam collimators reduce the over-ranging by collimating the x-ray beam at the start and end of the helical scanning.
Hybrid reconstruction technique (GE Healthcare):
Enables generation of images in the over-ranging portion of the x-ray beam to reduce extent of un-used x-ray beam.
Schilham et al. Radiographics 2010

## Slide 15: Choice of detectors, feed & pitch
Region of interest | Pitch | Detector configuration | Table feed
CT head helical | < 1 (overlapping) | Narrow | Slow
CT chest | ≥ 1 | Wide | Fast
CT chest (uncooperative) | Highest possible | Widest possible | Fastest possible
CT abdomen | ≥ 1 | Wide | Fast
CT abdomen (uncooperative) | Highest possible | Widest possible | Fastest possible
Small children (body CT) | Highest possible | Widest possible | Fastest possible
Choice of pitch and detector configuration can change based on
the MDCT scanner and required slice thickness.

## Slide 16: Summary
Detector configuration, pitch, and table feed are important scan parameters.
These parameters affect the scanning speed and radiation dose based on the CT vendor.
Users must use these parameters cleverly to meet the scanning requirements.