## Slide 1: Radiation Dose Management in Computed Tomography
L02
Overview of Scan Parameters
Material of the e-learning program
IAEA, 2017

## Slide 2: Objectives
Understand CT scan parameters
Know differences in nomenclature of scan parameters

## Slide 3: Indication for CT scan
From radiation dose perspective, no scan parameter or image reconstruction detail is as important as a justified clinical indication for CT scan

## Slide 4: There are a lot of scan parameters!
KVp
Pitch
Detector configuration
Scanner efficiency
Special reconstruction techniques (IRT)
Cardiac CT dose reduction EKG pulsing
  ECG controlled mA modulation
  Automatic pitch adaptation
Patient centering
Image filter
Scan length
Gantry rotation time
Reconstruction kernel
Table speed
mAs
Automatic exposure control
# scan series

## Slide 5: Scan parameters
Localizer radiograph
Helical/axial
Tube current
Tube voltage
AEC
Rotation time
Table feed/speed
Pitch
Scan length
Detector configuration
Reconstruction:
  Thickness
  Kernel/algorithm
  Technique (FBP/IRT)
Number of phases

## Slide 6: CT Localizer Radiograph
Localizer radiographs are images obtained by moving CT table through the scan field while the X ray tube is stationary:
  Antero-posterior projection: when X ray tube is positioned above the patient
  Postero-anterior projection: Position of tube is under the patient
  Lateral projection: when the tube is positioned on side of the patient
GE | Philips | Siemens | Toshiba
Scout | Surview | Topogram | Scanogram

## Slide 7: CT Localizer Radiograph
Used for deciding the start and end location of the scan range
Critical for determining appropriate tube current when using automatic exposure control (AEC) or automatic tube current modulation
When patients are not centred adequately for localizer radiography (too low/ high or too much to the side), incorrect estimation of tube current for AEC can result, which can inadvertently increase or decrease the dose

## Slide 8: Scanning Mode: Axial vs Helical
Axial mode: comprises of two alternating phases
  Data generation: Tube rotates around the stationary patient to acquire all projections at the specified location
  Patient positioning: Table moves to the next specified location while X ray tube is off (no data generated)
Also called “step and shoot” scanning
GE | Philips | Siemens | Toshiba
Axial | Axial | Sequence | Scan & View

## Slide 9: Scanning Mode: Axial
Axial scanning mode is still used commonly for
  Head CT
  Diffuse lung disease (for expiratory/ prone images)
  CT guided procedures to confirm needle position
  ECG gated cardiac CT with prospective triggering
GE | Philips | Siemens | Toshiba
Axial | Axial | Sequence | Scan & View

## Slide 10: Dose Reduction with Axial Mode
For multiple series exams
  E.g. check for loculated effusions
    Low dose prone images: Small scan length, 80 kVp, 50 mAs
  Other Eg. Biopsy, limited scanning, HRCT lungs
Standard dose supine
Prone low dose
MGH Boston

## Slide 11: Dose Reduction with Axial Mode
After lesion localization, reduce dose for CT guided Bx
  Axial acquisitions
  Reduce scan length and mA &/or kVp
Helical
Axial
Localizer images
Bx needle

## Slide 12: Axial mode over helical in HRCT Lungs
Inspiration Helical
DLP= 419 mGy.cm
Expiration AXIAL
DLP= 86 mGy.cm
Prone Axial
DLP= 43 mGy.cm
MGH Boston

## Slide 13: Scanning Mode: Axial vs Helical
GE | Philips | Siemens | Toshiba
Helical | Helical | Spiral | Helical
Helical or Spiral scanning mode: X ray tube is continuously ON while the table moves at constant speed.
Most body CT examinations are performed with helical scanning mode.

## Slide 14: Tube current (mA)
Represents number of electrons flowing through the cathode per unit time
Measured in milliamperes (mA)
The photons emitted (photon fluence) from the focal spot per unit time is proportional to mA

## Slide 15: Tube current (mA)
Tube current time product (mAs)
  Tube current x gantry rotation time in seconds
  Effective mAs or mAs/slice = mAs/pitch
Most commonly adjusted scan parameter to increase or decrease radiation dose:
  50% mA reduction results in 50% dose reduction#
  50% mA increase results in 50% dose increment#
# If all other scan parameters are held constant

## Slide 16: Tube current (mA)
Can be fixed over entire scan range (fixed mA)
Can be automatically adjusted over the scan range (automatic exposure control or AEC)
Different CT scanners have different mA limits
  Some can operate at much higher mAs than others.
GE | Philips | Siemens | Toshiba
mA | mAs/slice | Effective mAs | mA

## Slide 17: Tube potential (kV)
Represents potential difference between cathode and anode, which drives electrons across X ray tube
Unit is kilovoltage (kV)
Affects both the quality & quantity of generated X ray spectra
There is non-linear relationship between KV and dose
Dose changes by the square of kV change#
# If all other scan parameters are held constant

## Slide 18: Tube potential (kV)
Most scanners require users to specify a fixed kV for CT protocols
  Generally, smaller patients (particularly children) should be scanned at low kV (70-100 kV)
  Generally, CT angiography or contrast enhanced CT can be scanned at low kV (80-100 kV)
Some CT have automatic kV selection where CT scanner picks up the right kV:
  Care kV (Siemens)
  KV Assist (GE)

## Slide 19: Table Movement
Table Increment (mm)
  Term used for table travel in axial scanning mode
  Represents distance travelled by the table in one 360°rotation of tube
Table Feed (mm/rotation)
  Term used for table travel in helical scanning mode.
  Represents distance travelled by the table in one 360° rotation of tube
GE | Philips | Siemens | Toshiba
Speed | Table Speed | Table Feed | Couch Speed
GE | Philips | Siemens | Toshiba
Interval | Increment | Feed | Couch Movement

## Slide 20: Pitch
Beam Pitch
  Table feed (mm) divided by the width of collimated beam (mm)
Slice Pitch
  Table feed (mm) divided by selected detector collimation (mm)
  Detector collimation = width of each effective data channels x number of detector rows)
GE | Philips | Siemens | Toshiba
Pitch | Pitch | Pitch | CT Pitch factor

## Slide 21: Pitch: Considerations
Pitch is now a small factor in dose for modern MDCT
Scanners are dose efficient
Some CT change mA to compensate for pitch change
  Higher pitch: higher mA → similar dose → faster scanning#
  Lower pitch: lower mA → similar dose → slower scanning#
  Example: Siemens and Philips
Other CT partially change mA to compensate for pitch change
  Higher pitch: lower dose → faster scanning#
  Lower pitch: Higher dose → slower scanning#
  Example: GE and Toshiba
#If all other scan parameters are held constant

## Slide 22: Detector Arrays
Detector arrays can be fixed or variable in configuration.
Fixed detector array: Each detector row has same width
  Example: GE 64 slice CT
  64 detector rows each 0.625 mm thick
Variable detector array: Central rows are thinner than outer.
  Example: Siemens 16 slice CT
  16x0.75mm (central) with 4x1.5 mm on outer sides

## Slide 23: Detector Configuration
GE | Philips | Siemens | Toshiba
Detector configuration | Collimation N x T (mm) | Detector configuration | Detector configuration
4 x 1.5 16 x 0.75 4 x 1.5
Represents number of data channels along patient length (or z-axis) multiplied by the effective detector row thickness of individual data channel.

## Slide 24: Gantry Rotation Time
Time needed for x-ray tube and detector array to complete one 360° rotation
Different scanners have different minimum rotation times: 0.25 seconds is the lowest at the time of publishing this presentation
Faster gantry rotation: faster scanning
  Children
  Moving patients (uncooperative)
  Moving organs (heart)

## Slide 25: Scan Length
Dose increases with increase in scan length
Scanning must be restricted to what is essential
Extending scan range beyond region of interest increases dose
Localizer radiograph for CT abdomen
MGH Boston

## Slide 26: Scan length
95% of CT abdomen have extra images acquired beyond the area of interest.
Study showed that 1280 extra images were acquired in 106 patients equivalent to that in 12 -16 CT exams!
Limit scan to intended anatomic area to cut dose by 10%
  Abdomen-pelvis:
    Just above diaphragm: Inferior pubic symphysis
  Chest:
    Routine: Apex to adrenals
    Pulmonary embolism: Apex to lung bases
Kalra MK, et al. Radiology 2004
Campbell J, et al. AJR 2005

## Slide 27: Length of Overlap Scanning
Minimize overlap when scanning two contiguous body regions
Overlap scanning 100% (100/100) 100% (49/49)
Overlap scan range 7-18.5 cm 6-15.5 cm
Mean overlap scan 13.2 cm 9.7 cm
Total overlap scan 1 317 cm 476.5 cm
Namasivayam et al. ARRS 2006

## Slide 28: Slice Thickness and Interval
Slice Thickness (mm)
  Represents nominal width of reconstructed image in the longitudinal axis
Slice interval (mm)
  Represents distance between two consecutive reconstructed images
GE | Philips | Siemens | Toshiba
Thickness | Thickness | Slice | Image thickness
GE | Philips | Siemens | Toshiba
Interval | Increment | Position increment | Reconstruction interval

## Slide 29: Slice Thickness
Choice of slice thickness depends on the clinical indication and body region being scanned.
  Thin slice → higher noise → higher contrast → less partial volume artifact
    Suitable for high contrast and subtle abnormalities
  Thicker slice →less noise → more artifacts
    Suitable for low contrast and large patients

## Slide 30: Image Thickness and Lungs: Noise and Details
5 mm
3 mm
0.75 mm
3 mm
5 mm
Thinner sections: Details & sharpness
Thicker sections: Lower noise
150 mAs, 10 mGy
40 mAs 2.7 mGy
0.75 mm
Acquire thin images but read off the thick images to reduce noise in low dose images.
Notes: 30

## Slide 31: Image Thickness and Mediastinum: Noise
1 mm
2 mm
3 mm
4 mm
5 mm
Thicker sections = lower noise
Especially in soft tissues
150 mAs; 120 kV
CTDIvol = 10 mGy
Notes: 31

## Slide 32: Reconstruction kernel
Represents a feature on the scanner which influences the smoothness and sharpness of images in transverse plane
B20
B31
B50
B70
Lower noise
Poorer edge delineation
Better contrast
Higher noise level
Better edge delineation
Poor contrast
Standard
Details
Lung
GE | Philips | Siemens | Toshiba
Algorithm | Reconstruction filter | Kernel | Filter Convolution

## Slide 33: Reconstruction kernel and Lungs: Noise and Details
Soft
Smooth
Detail
Bone
2.5 mm; 10 Noise Index; CTDIvol = 10 mGy
2.5 mm; 40 Noise Index; CTDIvol = 2 mGy
Notes: 33

## Slide 34: Reconstruction kernel and Mediastinum: Noise
Soft
Smooth
Detail
Bone
2. 5 mm; 10 Noise Index CTDIvol: 14 mGy
2. 5 mm; 40 Noise Index CTDIvol: 2 mGy
Notes: 34

## Slide 35: Reconstruction techniques
Analytical image reconstruction (filtered back projection):
  Image reconstruction algorithm which back projects the sinogram to the 2D image domain, keeping few workable assumptions
  More prone to increased noise and artifacts at low dose.
  Fast way to reconstruct CT images
Iterative image reconstructions are newer techniques
  Incorporates better mathematical CT model and iterates to reduce inconsistencies in the image reconstruction
  Lower noise and artifacts at low dose.
  Generally slower than FBP

## Slide 36: Image reconstruction
Different angles-set of scan projections
(Raw or sinogram data)
Filtered back projection
Iterative
reconstruction
Back projection
Filtration: Approximations
Reconstructed image
Better defined
mathematical estimates
Reconstructed image
Multiple iteration
(undefined)
Initial step:
Data acquisition
Later steps:
Image reconstruction

## Slide 37: Iterative reconstruction techniques
Most IRT result in changes in image appearance compared to FBP
Most IRT come in different strength of noise reduction potential
Initial implementation of IRT should generally start at low strength of noise reduction
IRT does not reduce dose by itself but rather allow user to reduce dose compared to FBP

## Slide 38: Iterative reconstruction techniques
GE | Philips | Siemens | Toshiba
ASiR Veo ASiR-V | iDose IMR | IRIS Safire Admire | AIDR AIDR 3D
ASiR level is selected at 10% increment (10-100%) with increasing percentage associated with lower noise | iDose level can be set on a scale of 1-7 with increasing number suggesting lower noise | These techniques work on a five point scale (1-5) with increasing number associated with lower noise | Three settings (mild, standard, strong) represent increasing strength of noise reduction.

## Slide 39: IRT help reduce noise in low dose images
IRIS 200 mAs
FBP 200 mAs
FBP 100 mAs
IRIS 100 mAs
FBP 50 mAs
IRIS 50 mAs
IRIS (Siemens) reduces image noise at lower radiation dose as compared to FBP

## Slide 40: IRT help reduce noise in low dose images
200 mAs FBP
50 mAs FBP
50 mAs ASIR
50 mAs MBIR
ASIR and Veo (GE) reduce image noise at lower radiation dose as compared to FBP
Notes: 40

## Slide 41: Creating CT Protocols: Starting points
Age specific doses
children lower doses than adults
Region specific
chest < abdomen CT
Each region: At least 3 dose bins of indications
Routine: chest and abdomen
Multiphase: diffuse lung disease, liver/pancreas
Bin specifics
Divide protocols into two groups
Size specific dose adjustment: AEC, kV selection
Phase-specific adjustment: kV, mA, scan ranges
Classify protocols into body regions
Scan speed specific adjustment: pitch, table feed
Low dose: lung nodules and kidney stones

## Slide 42: Summary
Understanding scan parameters helps optimize radiation dose
There are differences in nomenclature of scan parameters between different CT vendors.