## Slide 1: Radiation Dose Management in Computed Tomography
L09
Key Aspects of Head CT
Material of the e-learning program
IAEA, 2017

## Slide 2: Objectives
Understand key protocols based on clinical indications for head CT
Understand scan parameters used for head CT
Address unique issues with head CT

## Slide 3: Why is head unique for CT?
Head CT is one of the most frequently performed CT
Remember that MRI has replaced head CT for several clinical indications
Skull base (densest bone in body) has propensity to cause beam hardening artifacts
Head anatomy changes rapidly from base to vault
Differences between grey and white matter in the brain represent extremely low contrast for CT.

## Slide 4: How is head unique for CT?
Head size does not change much with patient size unlike chest and abdomen
Certain clinical indications allow substantial dose reduction due to high inherent contrast
  Paranasal sinuses
  Facial and skull bones
  Craniostenosis
  Patency of ventricular shunts
Certain clinical indications have higher dose
  Stroke work up with CT
  CT angiography/venography of head and neck

## Slide 5: Developing CT protocols
Always pre-save scan protocols on the scanner
Head CT is often an urgent, time-sensitive exam
  Stroke work up with CT
  Head trauma
Time means brain parenchyma for stroke work up so fiddling with scan parameters and errors can delay CT scanning

## Slide 6: Basic of CT Dose
Justified: Ensuring that CT is the right test
ALARA: Adapting Dose to patient size or age
Interpretable: Tailoring CT for specific indications

## Slide 7: Protocols: Meet Basic Standards

## Slide 8: Make indication driven protocols
Indication based protocols are a must for dose optimization
Routine head
CT angiography head with or without neck
Perfusion CT
Temporal bone CT
CT orbits
Paranasal sinuses CT
Shunt patency

## Slide 9: Head CT: Minimize movement during scanning
Give instructions to minimize voluntary movements
Place in head rest to decrease motion
Coax children properly or use sedation
Uncooperative patients: Use faster scanning (helical scanning is a viable option in head)
Never perform more than one repeat acquisition CT if initial scan had motion
ONLY REPEAT ONCE if and when additional surety, sedation, or added legal restrain to decrease motion
  Faster scanning with faster gantry rotation, helical mode, higher pitch, and/or thicker beam collimation can help

## Slide 10: Head CT: Contrast Injection
Avoid suboptimal or non-diagnostic head CT from poor contrast injection technique
Make sure there is good intravenous (IV) access
Always flush to minimize risk of contrast extravasation
Inject at reasonable rate based on IV access
Always time the bolus for CT angiography
Avoid guessed delays or fixed delays
  Circulation times differ between patients
  Heart functions differ between patients

## Slide 11: Head CT: Good points
Localizer radiographs
  Make sure that head is well positioned
  Make sure that the head centering is good
  When possible head should be in head rest
  Use lower dose for localizer radiographs
For transverse CT images, specify
  Scan range
  Number of scan series
  Scan parameters for each series

## Slide 12: Axial versus Helical mode
Experience of MGH, Boston: Head CT performed with helical mode
  Advantage: Quick, MPR, no difference in quality
  Average CTDI vol = 45 - 60 mGy
  Disadvantage: No gantry tilt and higher eye dose
Many prefer – Axial scanning mode
  Advantage: lower dose (??), gantry tilt possible
  Disadvantage: Slower acquisition, motion artifacts
MGH= Massachusetts General Hospital
MPR= multiplanar reformats

## Slide 13: Axial versus helical CT
M. Straten, et al. AJNR 2007
Tan et al. AJNR 2009
McDermott, et al. Med. Phys. 2009
Studies suggest that thinly collimated helical brain CT is at least as good as thickly collimated axial CT
Some studies report eye dose from helical CT is greater than eye dose from titled axial scanning of the head
Other publications report no increase in eye dose with helical scanning in children as compared to axial mode

## Slide 14: Helical Mode for Head CT
There are special considerations to head CT performed in helical scanning mode
Overlapping pitch (~0.5): to decrease artifacts from rapidly changing anatomy
Thin beam collimation: To decrease radiation dose from over-ranging at the beginning and end of the helical acquisition
No gantry tilt allowed
Faster scanning compared to axial mode

## Slide 15: Head CT: Tube Current
Head CT can be performed with fixed mA or AEC
Prior studies have reported potential for reducing the fixed mA substantially for head CT
Substantial dose reductions have been reported with AEC
  % decrease depends on baseline protocol
AEC is also useful for neck CT
  Variations in size of shoulders
  Dramatic change in thickness from neck to shoulders requires benefit from change in mA with AEC
  Mullins et al. AJNR 2004
  Russell et al, AJNR 2008
  Smith et al. Radiology 2008

## Slide 16: Optimizing Dose: Tube potential
Traditionally routine adult head CT is performed at 120 kV
KV can also be selected as a fixed value or automatic kV selection technique.
CT angiography protocols benefit from lower kV (80-100kV)
  Smaller vessels are better seen
  Lower contrast volume or injection rate possible
  Lower radiation dose compared to higher kV
CT perfusion: Should be performed at 80 kV
  Maximizes sensitivity to detect iodine in the brain
  Helps reduce the dose dramatically compared to higher dose at 120 kV

## Slide 17: Good CT Exam: Other parameters
Scan series
  Must be minimum required
  When multiple- dose should not be multiple folds higher
Avoid rescanning same region
  E.g., head and temporal bone, face and sinuses
Scan length: Targeted and focused
Beam collimation: Lower is better to decrease dose for head
Remove extraneous hardware
  Brown, et al. AJR 1999

## Slide 18: Good CT Exam: Eye shielding?
Lead shield help reduce eye dose by 44% for head CT
  Can cause severe artifacts over the orbits
Bismuth in plane shields also reduce radiation dose to eyes with less severe artifacts
  40-53% dose reduction to lens with Bismuth shields
  Controversial due to effect on image quality
Ngaile, J. E. et al. Radiat Prot Dosimetry 2008 130:490-498;
Hopper et al. AJNR 2001

## Slide 19: Points against shielding
Heany et al. Australas Phys Eng Sci Med. 2006
  Angulation is superior to Bismuth eye shields
    88% reduction versus 48% reduction
Leswick et al. Radiology 2008
  AEC is better than Bismuth thyroid shields
    78% dose reduction versus 42%
    Other publications support use of Bismuth

## Slide 20: Pediatric head CT
Perform only when appropriate, use brain MRI to avoid need for head CT when possible.
Lower dose head CT should be used in children compared to adult head CT.
Create specific protocol for ventricular shunt patency: low kV and fixed low mA
Create separate protocol for Craniostenosis: Low KV (80 kV) and low mAs (ex. 50 mAs)

## Slide 21: Craniostenosis: 80 kV, 30 mAs, helical pitch 1.4:1, 0.04 mSv

## Slide 22: Post trauma: 120 kV, 90-140mA, p = 0.984, 5 NI, 5 mm – 2.5 mm

## Slide 23: Sinus and Orbits
Facial bones and sinuses can be scanned at extremely low radiation dose
  High tissue contrast between bones and air offers scope for reduced radiation dose CT
  100-120 kV and 40-50 mAs generally sufficient for sinuses
  Helical scanning allows isotropic image reconstruction in other planes without need to acquire additional series in coronal plane

## Slide 24: Stroke work up CT
Doses with comprehensive stroke protocol CT can be as high as 16-20 mSv
  Non-contrast CT
  CT angiography (CTA)
  CT perfusion
  Post contrast images of brain
For CTA: use lower kV (≤ 100 kV)
For CT perfusion: use 80 kV

## Slide 25
Iterative Reconstruction in Head CT
Estimated dose:
CTDIvol: 25 mGy
DLP: 459 mGy.cm
Estimated dose:
CTDIvol: 66 mGy
DLP: 1224 mGy.cm
Prior scan: FBP
Follow up: ASIR
60 % dose reduction

## Slide 26: Iterative reconstruction in Head CT
FBP 150 mA
FBP 150 mA
FBP 75 mA
75 mA-S3
75 mA-S2
75 mA-S1
IR results in lower noise in reduced dose images

## Slide 27: Summary
Head CT protocols must be organized based on clinical indications
Care must be taken to reduce motion artifacts
CT angiography and CT perfusion must be performed with low kV
Pediatric head CT should be performed at low radiation dose compared to adult head CT