## Slide 1: Radiation Dose Management in Computed Tomography
L06
CT Dose Metrics and Tracking
Material of the e-learning program
IAEA, 2017

## Slide 2: Objectives
To understand basic CT dose metrics:
  CTDIvol
  DLP
  SSDE
To understand advantages and logistics of different dose tracking options for CT

## Slide 3: Radiation Dose Measurements
Radiography-
Entrance skin dose (ESD)
CT- CTDIvol, DLP
Dose comparisons between radiography and CT in terms of skin dose are not appropriate.

## Slide 4: Radiation Dose Measurements
Two standardized dose metrics displayed on modern day CT scanners
  CT dose index volume (CTDIvol)
  Dose length product (DLP)
Size-specific dose estimates (SSDE) have been proposed

## Slide 5: CTDIvol
CTDIvol is a standardized method to measure radiation output from the scanner
  These dose measurements are obtained from 16 or 32 cm diameter acrylic phantoms at CT factories
    16 cm Head CT
    32 cm Body CT
  Phantom size used for its calculation is displayed on CT
  Measured in milli Gray (mGy)
  CTDIvol does not represent actual patient doses
  Represents average dose within each slice

## Slide 6: Dose Length Product (DLP)
DLP is product of CTDIvol (mGy) and length of scanned area (cm)
  DLP is also calculated by the scanner based on CTDIvol
  Represents integrated dose over the entire scanned area
  Unit is mGy.cm
  Again, DLP does not represent patient-specific dose

## Slide 7: Current Dose Displays on CT
Provides information on amount of dose from the exam
Useful to track across patients and protocols: Registry and Audit
Compare and improve protocols across institutions.
Regulatory compliance
Phantom dose only (16 or 32 cm only)
No patient bod size compensation
  Underestimates for small patients
  Overestimates for large patients
Does not account body composition
The Good!
The Bad!
WANTED: Preferably Viable!
Better dose descriptor to account for size, and body region of the patient

## Slide 8: Size-Specific Dose Estimates (SSDE)
American Association of Physicists in Medicine (AAPM) report # 204 introduced SSDE for CT
Helps convert CTDIvol into patient size specific doses using conversion factor derived from
  Antero-posterior (AP) width
  Lateral width
  Sum (AP + lateral) width
  Effective diameter (square root of (AP x lateral width)

## Slide 9: Size-Specific Dose Estimates (SSDE)
Extends application of CTDIvol to patient size beyond 16 and 32 cm
Still does not account for patient composition, like CTDIvol
Still does not represent actual organ doses, like CTDIvol
Should not be used to estimate specific patient doses

## Slide 10: Current Post- scan Dose Display
Legacy (really old) CT
  No automatic post-scan dose capture!!
Legacy (not that old) CT
  Dose information page
    Looks different for different CT vendors
    Image capture of mathematical data!!
Newer CT systems
  DICOM Dose Structured Reports capture doses and scan parameters in an downloadable, analyzable format
  Dose information page (sticks around!)

## Slide 11: Sample Dose Page: Example A
Most parameters (kVp, mAs, pitch) not displayed

## Slide 12: Sample Dose Page : Example B
Some parameters (scan length, pitch) not displayed

## Slide 13: How scan parameters affect CT doses?
Scan parameters | Effect on CTDIvol &/or DLP#
Axial scan mode | No over-ranging (extra radiation at the start and end of scanned volume)
Helical scan mode | Over-ranging present at start and end of scanned volume
Tube current | Direct relationship mA ∝ CTDIvol
Tube potential | Stronger direct relationship Dose ∝ (kV)2
# If all other scan parameters are held constant

## Slide 14: How scan parameters affect CT doses?
Scan parameters | Effect on Dose#
Pitch | Some scanners do not allow change in dose with change in pitch Other CT scanners, CTDIvol decreases with increasing pitch and vice versa
Gantry rotation time | Relationship identical as mA and dose. Direct linear relationship with CTDIvol
Scan length | Does not affect CTDIvol Direct linear relationship between scan length and DLP
# If all other scan parameters are held constant

## Slide 15: More Dose terms: Diagnostic Reference Levels (DRLs)
DRL is a reference dose for identifying level of radiation dose for a diagnostic x-ray exam
  Typically at 75th percentile of observed doses from survey of facilities (doses at 75% facilities are below DRL)
  Different countries and organizations have different DRLs for same CT protocols
DRLs help in CT radiation dose optimization efforts
DRLs are for average size patients and are often higher for smaller patients and lower for large patients

## Slide 16: More Dose terms: Notification and Alert Values
Notification value: Dose level of a series or scan that causes a pop up warning notice to be displayed prior to scanning.
  AAPM recommends values of 80 mGy (adult head), 50 mGy (adult torso)
Alert value: Cumulative dose level of entire study (including both done and being done) that results in a pop up alert to be displayed on the scanner.
  US FDA recommends alert value of 1 Gy
Both values help radiographers in case of exceeding specified dose levels.
National Electric Manufacturers Association (NEMA) XR 25 standard

## Slide 17: Dose Tracking: Why bother?
Avoid new or duplicate CT exams
Indicate need for scan parameter adjustment versus other scanners, protocols or facilities
Know the outcome of optimization efforts
Make judgment of need in patients who have already have multiple prior CT exams
Check compliance with recommended DRLs

## Slide 18: Dose Tracking methods
Working without computer program to record dose
Manual writing or manual entry of dose metrics
    Error prone
    Time consuming
    Tedious: what to save and what to let go?
    Complex: for multi-part or multi-series exams
    Privacy concerns: Lost paper, corrupted files, malware

## Slide 19: Dose Tracking methods
Automatic extraction of dose metrics is better than manual capture
Reliable
Accurate
Automatic database creation
    Only facility level dose audits
    Set up expertise may be needed
    Ongoing maintenance

## Slide 20: Dose Tracking: Freeware
RADIANCE- Radiation Dose Intelligent Analytics for CT examinations (http://www.radiancedose.com)
Dose Utility
DoseRetriever
GROK (Generalized Radiation Oncology Toolkit in PixelMed Library)
OpenREM.org (http://openrem.org)

## Slide 21: Commercial Dose Monitoring
Valkyrie (Columbia University)
PENMET (Patient Exposure Management NETwork) from Clinical miscrosystems, Inc. (http://www.penmet.com)
DoseMonitor (PHS Technologies Group LLC Bracco) (www.dosemonitor.com)
DoseMetrix from Primordial Inc. (http://www.primordialdesign.com/home.html)
Radimetrics Bayer HealthCare (http://www.medrad.com/enus/info/products/Pages/Radimetrics - Enterprise - Platform.aspx)
DoseWatch (GE) (www.3.gehealthcare.com/en/products/dose_management/dosewatch)
RightDose (Siemens) (http://www.healthcare.siemens.com/medical - imaging/lowdose/)
Dose Track (Spectra) (https://www.sectra.com/medical/dose_monitoring)

## Slide 22: National and International Dose Tracking
In the United States, the American College of Radiology Dose Index registry (ACR DIR) automatically retrieves and archives doses on millions of CT examinations for benchmarking and quality assurance
Elsewhere, in countries like UK, Germany, Finland, and others there are additional dose tracking registry

## Slide 23: Dose Tracking: Should we do it?
Yes:
  Ensure that we do no harm (primum non nocere)
  Ensure that our CT doses are with in DRLs
  Identify and eliminate rogue protocols or practices
  Assess results of dose optimization initiatives
  Identify patients and protocols exceeding alert values
  Assess need for equipment upgrade or change

## Slide 24: Summary
Dose metrics of CT are different from other modalities.
Effect of change in scan parameters on dose metrics are important to know.
Remember and use the DRLs initially, they are good starting points for dose optimization.