## Slide 1: Radiation Dose Management in Computed Tomography
L03
Tube Current and Automatic Exposure Control

## Slide 2: Objectives
Understanding tube current
Basis of AEC technique
Types of AEC techniques
Application of AEC for dose reduction in body regions
Limitations of AEC techniques

## Slide 3: Tube current (mA)
Represents number of electrons flowing through the X-ray tube per unit time
Measured in milliamperes (mA)
The photons emitted (photon fluence) from the focal spot per unit time is proportional to mA
Tube current time product (mAs)
  Tube current x gantry rotation time in seconds
  Effective mAs or mAs/slice = mAs/pitch
Most commonly adjusted scan parameter to increase or decrease radiation dose.
  50% mA reduction results in 50% dose reduction#
  50% mA increase results in 50% dose increment#
# If all other scan parameters are held constant

## Slide 4: Tube current (mA)
Can be fixed over entire scan range (fixed mA)
Can be automatically adjusted over the scan range (automatic exposure control or AEC)
Different CT scanners have different mA limits
  Some can operate at much higher mAs than others.
GE | Philips | Siemens | Toshiba
mA | mAs/slice | Effective mAs | mA

## Slide 5: Fixed Tube Current
May be the only option in certain legacy CT systems without AEC.
For appropriate use of fixed mA
  Adjust fixed mA for different clinical indication
  Create size-specific fixed mA settings for each indication
  Pediatric CT will require multiple weight or size group fixed mA protocols

## Slide 6: Fixed mA
z-axis
mA per slice
Attenuation per slice
200 mA
200 mA
200 mA
200 mA
200 mA
200 mA
200 mA
200 mA
Fixed mA
Kalra et al. RSNA 2005
There is no change in applied mA with fixed mA technique

## Slide 7: Automatic Exposure Control (AEC)
Kalra et al. Radiology 2004
Defined as a set of techniques that enable automatic adjustment of the tube current in the x-y plane (angular modulation), along the z-axis (z-axis modulation), or both (combined modulation) according to the size and attenuation characteristics of the body part being scanned and achieve constant CT image quality with lower radiation dose.
Analogous to the automatic exposure-control or photograph-timing techniques used in conventional radiography.
Adjust mA to better indicators of patient size than weight and cross-sectional diameters
These are generally attenuation profile:
  Patient cross-sectional dimension
  Attenuation or “density”
Word “automatic” is a “misconception”
AEC requires important user input to operate optimally
AEC changes mA to selected slice thickness, pitch, kVp

## Slide 8: Projections and Positions
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
Z-axis section positions
Section position 9
P= Projection angles at
a given section position
P 1
P2
P 3
P 4
Y-axis
X-axis
Kalra et al. Radiology 2005
Angular projection positions

## Slide 9: Automatic Exposure Control (AEC)
Maintains image quality (noise) by adapting current to changing body size and attenuation.
  Slice-to-slice (longitudinal or z-axis modulation)
  Within each slice (angular or x-y modulation)
  Both: z- and angular- (xyz/combined modulation)
  ECG modulated tube current or ECG pulsing

## Slide 10: AEC Techniques
Vendors | Z-axis | Angular | Combined
GE | Auto mA | Smart mA | Auto mA 3D
Philips | Z - DOM | Dose right dose Modulation (DOM) | -
Siemens | Z-axis exposure control (ZEC) | CARE Dose | CARE Dose 4D
Toshiba | Real Exposure Control (Real EC) | - | SURE Exposure 3D

## Slide 11: Fixed mA
z-axis
mA per slice
Attenuation per slice
200 mA
200 mA
200 mA
200 mA
200 mA
200 mA
200 mA
200 mA
Fixed mA
Kalra et al. RSNA 2005
There is no change in applied mA with fixed mA technique

## Slide 12: Angular AEC or x-y modulation
54 mA
54 mA
200 mA
200 mA
146 mA
146 mA
146 mA
146 mA
Localizer Information
Kalra et al. RSNA 2005
Angular modulation may be based entirely on information from localizer radiograph or partly from actual scanning (online technique).
Angular AEC involves change in mA at different angular projections

## Slide 13: Angular Modulation
In Practice:
User specifies mA or effective mAs (mAs/pitch)
Scanner modulates current for each slice
  Real-time (CARE Dose, Siemens; DOM, Philips)
  Scout image (Smart mA, GE; Sure Exposure, Toshiba)
Limitation:
  Appropriate mA or effective mAs is pre-requisite to dose optimization
Greess et al. European Radiology 2004

## Slide 14: Angular Modulation
CARE Dose
DOM
Kalra et al. Radiology 2004

## Slide 15: Angular AEC
Angular AEC = 159 mAs
22% Dose Reduction
Fixed Current = 200 mAs
SD = 12.0
  10 - 60% dose reduction (mean reduction 22%) compared to fixed mA without loss of image quality
Kalra et al. Radiology 2004
SD = 12.3

## Slide 16: Longitudinal modulation (z-AEC)
User specifies image quality parameter based on required image quality
Different CT vendors use different nomenclature for image quality parameters
Information from localizer radiograph (1 or both) used to adapt mA
Limitations: Users decide desired noise or image quality for different clinical indications
Kalra et al. Radiology 2004

## Slide 17: Longitudinal modulation (z-AEC)
mA per section position
Attenuation per section position
162 mA
162 mA
162 mA
162 mA
208 mA
208 mA
208 mA
208 mA
Fixed mA Z-axis AEC
Chest – lower x-ray attenuation
Abdomen – greater x-ray attenuation
z-AEC involves change in mA at different slice positions
z-axis

## Slide 18
 | Parameter | Description
Auto mA GE | Noise index Minimum-Maximum mA | User defined noise in overall image User defined mA range for achieving noise index
Z-DOM Philips | Baseline mAs | Used as reference to obtain constant image quality along the z-axis
CareDose 4D* Siemens | Quality reference mAs | Means image quality similar to that obtained with specified effective mAs in a standard size adult
Sure Exposure 3D* Toshiba | Standard deviation | Implies user requires images with specified noise (standard deviation)
(* Imply combined modulation technique)
Kalra et al. Radiology 2004
Longitudinal modulation (z-AEC)

## Slide 19
Combined Modulation
160 mA
120 mA
120 mA
124 mA
165 mA
124 mA
z-axis
160 mA
165 mA
Combined modulation (angular and longitudinal) combines the two AEC techniques

## Slide 20
Combines angular and longitudinal modulation techniques
Application is similar to longitudinal modulation technique
Users have to specify the required image quality metric like in longitudinal modulation
May provide additional radiation dose reduction as compared to separate use of angular and longitudinal modulation techniques
Combined Modulation

## Slide 21: Auto mA and Auto mA 3D (GE)
Auto mA = longitudinal modulation
Auto mA 3D = Longitudinal modulation + Smart mA (x-y)
Image quality parameter: Noise Index
User must specify required noise index for application
Noise index is noise (SD) in image of uniform phantom.
  Lower noise index = less noise = higher dose
  Higher noise index = more noise = lower dose

## Slide 22: Auto mA and Auto mA 3D (GE)
Choice of noise index depends on clinical indication
  Routine abdomen needs lower noise index (better quality) compared to higher noise index (lower quality- lower dose) for kidney stone CT
  Likewise, routine chest CT (higher tissue contrast) needs higher noise index compared to routine abdominal CT (lower tissue contrast)
Minimum and maximum mA are needed to control the extent of tube current change.
Uses the last acquired localizer radiograph to determine patient size for estimation of tube current.

## Slide 23: Auto mA and Auto mA 3D (GE)
Like other AEC techniques, it is profoundly affected by off-centering of the patient.
Noise index is not independent of prospective slice thickness
  Allow increase in noise index with decrease in slice thickness to keep constant dose compared to thick slices.

## Slide 24: Auto mA and Auto mA 3D (GE)
Noise index 10/2.5 mm
50% less dose
Noise index 15/ 2.5mm
70% less dose
Noise index 12.5/5mm
26% less dose
Noise index 15/2.5 mm
70% less dose
84 mAs
34 mAs
37 mAs
10 mAs

## Slide 25: Longitudinal modulation (Toshiba)
The system offers four levels of image noise (standard deviation) to match diagnostic needs.
Real EC, Toshiba

## Slide 26: Real EC and Sure Exposure 3D (Toshiba)
Standard deviation (SD) is image quality metric
SD behaves similar to noise index
  Higher SD implies lower mAs and lower dose
  Lower SD means higher mAs and dose
Choice of SD should again be based on clinical indication for the scan
  High contrast regions (chest) and indications (stones, pulmonary nodules) can be scanned at higher SD compared to lower tissue contrast regions and indications

## Slide 27: ZEC and CARE Dose 4D (Siemens)
Quality reference mAs
200
Requires user to suggest quality reference mAs (QRM)

## Slide 28: ZEC and CARE Dose 4D (Siemens)
QRM implies average effective mAs for a reference patient
Reference patient is defined as
  Typical adult with weight of 70-80 kg (for adult CT)
  Typical child with weight of 20 kg (for pediatric CT)
Instead of specifying the mA range, this technique allows users to control strength of mAs change for different size patients (Slim or Obese)
  For slim: weak decrease, average decrease and strong decrease settings control extent of mAs decrease
  For obese: weak increase, average increase, strong increase settings control extent of mAs increase

## Slide 29: CARE Dose 4D, Siemens
Fixed current = 20 mGy
Care Dose 4D = 8.8 mGy
Care Dose = 15 mGy

## Slide 30: ECG Based Tube Current Modulation
50
400
mA
20% - 50% dose reduction
Decreases radiation dose by 20 - 50% depending on heart rate by performing temporal modulation of mA based on ECG compared to fixed mA throughout cardiac cycle
mA can be decrease during certain phases of R-R interval as desired by the users
Extent of decrease in mAs is about 5 - 20% of the baseline mAs

## Slide 31: Practical implementation of AEC
Several studies have reported substantial dose reduction with use of AEC without compromising desired image quality
Most body CT exams must be performed with AEC
AEC techniques will change the mAs based on patient size
  For “x” image quality, a smaller patient will get lower dose than a larger patient
  Extent of change of radiation with AEC can be changed by changing image quality
  Strength of change in mAs with AEC can also be controlled

## Slide 32: Caveats of AEC: Clinical Indication
AEC techniques recognize the patient size to adapt mAs and radiation dose
AEC does not recognize the clinical indication!
Users must change image quality to adjust dose according to clinical indications
  Routine abdomen (higher quality) than kidney stone
  Routine chest (higher quality) than lung nodule follow up CT
Bottom line: Make indication specific CT protocols

## Slide 33: Caveats of AEC: Centering of patient
Importance of adequate centering of the patient in the gantry is paramount to optimal AEC operation
  Height of the table
    Avoid too low or too high heights to avoid off-centering
  Side to side centering
    Avoid placing patient to one side of the table
Off-centering can lead to severe artifacts and mis-calculation of required mAs with AEC

## Slide 34: X-ray filtration
  Decrease “soft x-rays” that constitute absorbed radiation but never reach detectors
  Efficient x-ray filters selectively remove these soft rays and decrease absorbed dose
  Bow-tie or beam shaping filters reduce surface radiation dose by minimizing dose in thinner portions of patient anatomy
Kalra et al. Radiology 2005

## Slide 35: Patient centering with in bow-tie filter
Uniform x-ray
at detector
Bow-tie filter
Detectors
Patient
X-ray tube
X-rays
Li et al. AJR 2007

## Slide 36: Patient off-centered in gantry isocenter
Uniform x-ray
at detector
Bow-tie filter
Detectors
Patient
X-ray tube
X-rays
Li et al. AJR 2007

## Slide 37: AEC caveats: Position of arms
AEC techniques can not recognize arms from other body regions (chest or abdomen)
When performing chest and abdomen CT
  Arms placed by the side of body for localizer and CT: mAs and dose increased with AEC
  Arms placed by the side for localizer but raised for CT: mAs and dose increased with AEC
  Arms raised above head for localizer and by side for CT: Lower mAs applied with AEC than needed

## Slide 38: AEC can not recognize arms from the body in the localizer
20 mGy
14 mGy
20 sec later with
no change in scan factors
Kalra et al. AJR 2004

## Slide 39: AEC caveats: Metallic prosthesis
Certain AEC techniques do not recognize metallic implants from body tissues in the localizer radiograph
  These techniques increase radiation dose in region of large metallic prosthesis without change in severity of artifacts
Other techniques recognize and exclude metallic implants when estimating mAs from the localizer

## Slide 40: Longitudinal modulation: Metallic implants
Fixed mAs 120
Z-modulation 47 mAs
Z-modulation 150 mAs
Rizzo et al. AJR 2004; Dalal et al. Radiology 2005

## Slide 41: AEC caveats: Image Quality Requirements
AEC techniques require users to specify desired image quality as a starting point for dose modulation
Selection of image quality for different clinical indications is not easy
  Review of literature helps identify some guidelines
  Gradual change and trial helps set appropriate levels

## Slide 42: AEC: Summary
Tube current is the most important scan parameter for adjusting radiation dose
AEC techniques help reduce dose compared to fixed mAs
When using AEC, users must remember caveats and pitfalls of AEC